International Union of Geological Sciences

Manual of Standard Methods

for

Establishing the Global Geochemical Reference Network

Chapter 6

Geoanalytical Methods and Requirements

Manfred Birke1,5, Gwendy E.M. Hall2,5, George Morris3,5,

Anna Ladenberger3,5, Alecos Demetriades4,5

1 Bundesanstalt für Geowissenschaften und Rohstoffe, Hannover, Germany

2 Geological Survey of Canada, Ottawa, Ontario, Canada

3 Geological Survey of Sweden, Uppsala, Sweden

4 Institute of Geology and Mineral Exploration, Athens, Hellenic Republic

5 IUGS Commission on Global Geochemical Baselines

2022

Published by

The International Union of Geological Sciences

Commission on Global Geochemical Baselines

It is recommended that reference to this part of the Manual should be made in the following way:

Birke, M., Hall, G.E.M., Morris, G., Ladenberger, A. & Demetriades, A., 2022. Geoanalytical Methods and Requirements. Chapter 6 In: Demetriades, A., Johnson, C.C., Smith, D.B., Ladenberger, A., Adánez Sanjuan, P., Argyraki, A., Stouraiti, C., Caritat, P. de, Knights, K.V., Prieto Rincón, G. & Simubali, G.N.  (Editors), International Union of Geological Sciences Manual of Standard Methods for Establishing the Global Geochemical Reference Network. IUGS Commission on Global Geochemical Baselines, Athens, Hellenic Republic, Special Publication, 2, 335−385.

Published by

The International Union of Geological Sciences

Commission on Global Geochemical Baselines

P.O. Box 640 37, Zografou

GR-157 10 Athens

Hellenic Republic

http://www.iugs.org/

http://www.globalgeochemicalbaselines.eu/

National Library of Greece Cataloguing in Publication Data

A catalogue record for this electronic book is available from the

National Library of Greece

ISBN: 978-618-85049-1-2

Contents

6.1. Introduction	339

6.2. Global Geochemical Reference Network analytical requirements	343

6.2.1. Use of standard reference materials	344

6.2.2. Proposed arrangements for reference sample analysis	345

6.2.3. National and international geoanalytical requirements	346

6.2.4. Need for independent quality control	347

6.3. Analytical schemes used in multinational continental-scale projects	348

6.3.1. Participating laboratories and analytical techniques	348

6.3.2. Sample preparation	349

6.3.3. Analytical methods used in the European Global Terrestrial Network	349

6.3.3.1. Analysis of stream water samples	349

6.3.3.1.1. Determination of anions	349

6.3.3.1.2. Determination of dissolved organic carbon	350

6.3.3.1.3. Determination of major cations, metals and other inorganic elements	350

6.3.3.1.4. Determination of mercury	350

6.3.3.1.5. Selection of stream water data	353

6.3.3.2. Analysis of stream and floodplain sediment samples	353

6.3.3.2.1. Determination of TOC	353

6.3.3.2.2. Determination of mercury	354

6.3.3.2.3. Total multielement determination by XRF	354

6.3.3.2.4. Total multielement determination by ICP-MS	354

6.3.3.2.5. Acid leachable multielement determination by ICP-AES	355

6.3.3.2.6. Selection of stream and floodplain sediment data	355

6.3.3.3. Analysis of residual soil samples	355

6.3.3.3.1. Determination of particle size distribution	355

6.3.3.3.2. Determination of TOC	355

6.3.3.3.3. Determination of mercury	356

6.3.3.3.4. Total multielement determination by XRF	356

6.3.3.3.5. Total multielement determination by ICP-MS	356

6.3.3.3.6. Acid leachable multielement determination by ICP-AES	357

6.3.3.3.7. Selection of residual soil data	357

6.3.3.4. Analysis of humus samples	357

6.3.3.4.1. Determination of mercury	357

6.3.3.4.2. Total multielement determination by ICP-MS	357

6.4.3.4.3. Selection of humus data	358

6.3.4. Quality assurance and control	358

6.3.5. Additional methods used in European continental-scale projects	359

6.3.5.1. Analytical methods used in the European Groundwater Geochemistry project	359

6.3.5.1.1. ICP-QMS analyses	359

6.3.5.1.2. ICP-AES analyses	362

6.3.5.1.3. Ion chromatography (IC)	363

6.3.5.1.4. Photometric analyses	363

6.3.5.1.5. Titration method	364

6.3.5.1.6. Atomic fluorescence spectrometry (AFS)	364

6.3.5.2. Analytical methods used in the European Geochemical Mapping of Agricultural and grazing land Soil project	364

6.3.5.2.1. Total element concentrations by XRF	364

6.3.5.2.2. Hot aqua regia extraction and determination by ICP-AES & ICP-MS	365

6.4. Analysis of rock samples	368

6.5. Determination of mineralogy by XRD	370

6.6. Particle size analysis	371

6.7. Summary of recommendations	372

References	375

Blank back page

## 6.1. Introduction

Progress in geochemical mapping, since it began in the 1950s, has been due principally to the development of new and improved instrumental techniques, providing better sensitivities for more elements more rapidly. Analytical requirements for the compilation of a global geochemical database, designed to facilitate firstly the establishment of the Global Geochemical Reference Network and secondly international geochemical mapping, need the best of the existing analytical techniques. Because the data are required for long-term reference purposes and comparisons, some special considerations need attention that should be attended to by the project’s Laboratory Committee.

Most geochemical mapping has been undertaken to assist in discovering mineral deposits.  Geochemical surveys carried out in the 1950s were considered a new method of prospecting to enable the mining industry to locate unexposed mineralisation, e.g., Cu, Pb-Zn or Ni, usually within a geographically restricted area. The work was, therefore, narrowly focused and, at best, semi-quantitative analytical techniques were used (Lakin et al., 1952; Ward, 1963; Stanton 1966; Levinson, 1974, 1980). National mapping programmes, which began 15 to 20 years later, expanded the approach started by the mining industry to provide information about a wider range of elements of possible economic interest (see Table 2-2, p.24-25, in Darnley et al., 1995). Since the mid-1970s national regional geochemical atlas programmes have begun to add elements primarily because of their environmental significance (e.g., Webb et al., 1978; BGS, 1978, 1991, 1992, 1993, 1999, 2000; Thalmann, 1984; Fauth et al., 1985; Xie et al., 1985; Lahermo et al., 1990; Koljonen, 1992; Lis and Pasieczna, 1995; Kadünas et al., 1999; Ottesen et al., 2000, 2010; Halamić and Miko, 2009; Caritat and Cooper, 2011; Xie et al., 2012; Andersson et al., 2014; Smith et al., 2014).

In general, however, the elements determined, and the analytical methods selected, have been chosen according to the presumed principal mineral potential of the region under investigation. Emphasis on mineral exploration considerations has influenced the types of samples collected, sample preparation, the analytical methods employed and the way they have been used. Cost-effectiveness, judged from a mineral exploration viewpoint, has been a prime consideration, and the quality of obtained geochemical data was not always a priority. This has not prevented large and unique data sets with important environmental implications from being acquired.

For both industry and national geological surveys, the cost was often and still is a limiting factor in the design of projects with the necessity to counterbalance sampling density, the comprehensiveness and quality of the data to be obtained, against the size of the area to be explored. This drawback has been overcome with the availability of low-cost multi-element analytical techniques, as has been shown in continental-scale geochemical projects carried out in Australia, China, Europe and the United States of America (Table 6.1). The data are still not, however, compatible because of different sampling, sample preparation and analytical techniques.

Table 6.1. Continental-scale projects carried out in Australia, China, Europe and the United States of America (USA) with project acronyms, sampling media collected, and analytical methods used.

| Country | Australia1 | China1 | Europe2 |  |  |  | USA1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Project | NGSA | CGB | FOREGS |  |  | GEMAS | NASGLP |
| Sampling medium | Sediment/Soil3 | Sediment/Soil3 | Soil4 | Sediment5 | Humus | Soil6 | Soil7 |
| Element |  |  |  |  |  |  |  |
| Ag | TOT, AR, MMI | DCA | TOT | / | / | AR, MMI | MA |
| Al | XRF, AR, MMI | WD-XRF | XRF | WD-XRF | / | WD-XRF, AR, MMI | MA |
| As | TOT, AR, MMI | / | TOT, AR | WD-XRF, AR | / | WD-XRF, AR, MMI | AAS |
| Au | FA, AR, MMI | AR | / | / | / | AR, MMI | / |
| B | AR | DCA | / | / | / | AR | / |
| Ba | TOT, AR, MMI | XRF | WD-XRF, AR | WD-XRF, AR | TOT | WD-XRF, AR, MMI | MA |
| Be | TOT, AR | MA | TOT | TOT | / | AR | MA |
| Bi | TOT, AR, MMI | MA | TOT | / | / | WD-XRF, AR, MMI | MA |
| Br | / | WD-XRF | / | / | / | WD-XRF | / |
| C | / | COMB | / | / | / | EA | / |
| Ca | XRF, AR, MMI | XRF | WD-XRF | TOT | / | WD-XRF, AR, MMI | MA |
| Cd | TOT, AR, MMI | / | TOT | TOT | TOT | AR, MMI | MA |
| Ce | TOT, AR, MMI | / | TOT | TOT | / | XRF, AR, MMI | MA |
| Cl | XRF | XRF | / | / | / | WD-XRF | / |
| Co | TOT, AR, MMI | MA | TOT, AR | TOT, AR | TOT | WD-XRF, AR, MMI | MA |
| Cr | TOT, AR, MMI | / | WD-XRF, AR | WD-XRF, AR | / | WD-XRF, AR, MMI | MA |
| Cs | TOT, AR, MMI | MA | TOT | TOT | / | WD- XRF, AR, MMI | MA |
| Cu | TOT, AR, MMI | MA | TOT, AR | TOT, AR | TOT | WD-XRF, AR, MMI | MA |
| Dy | TOT, AR, MMI | / | TOT | TOT | / | MMI | / |
| Er | TOT, AR, MMI | / | TOT | TOT | / | MMI | / |
| Eu | TOT, AR, MMI | / | TOT | TOT | / | MMI | / |
| F | ISE | / | / | / | / | WD-XRF | / |
| Fe | XRF, AR, MMI, TIT | XRF, MA, CAL | WD-XRF, AR | WD-XRF, AR | / | WD-XRF, AR, MMI | MA |
| Ga | TOT, AR, MMI | XRF | TOT | WD-XRF | TOT | WD-XRF, AR, MMI | MA |
| Gd | TOT, AR, MMI | / | TOT | TOT | / | MMI | / |
| Ge | TOT, AR | MA | / | / | / | AR | / |
| Hf | TOT, AR | MA | TOT | TOT | / | WD-XRF, AR | / |
| Hg | AR, MMI | / | TOT | TOT | TOT | AR, MMI | AR |
| Ho | TOT, AR | / | TOT | TOT | / | / | / |
| I | / | AFC | TOT | / | / | ED-XRF | / |
| In | AR | MA | TOT | / | / | AR, MMI | MA |
| K | XRF, AR, MMI | XRF | / | WD-XRF | / | WD-XRF, AR, MMI | MA |
| La | TOT, AR, MMI | / | TOT | TOT | TOT | WD-XRF, AR, MMI | MA |
| Li | AR, MMI | MA | / | TOT | / | AR, MMI | MA |
| Lu | TOT, AR | / | TOT | TOT | / | / | / |
| Mg | XRF, AR, MMI | MA | WD-XRF | WD-XRF | / | WD-XRF, AR, MMI | MA |
| Mn | XRF, AR, MMI | XRF | WD-XRF, AR | WD-XRF, AR | / | WD-XRF, AR, MMI | MA |
| Mo | TOT, AR, MMI | / | TOT | TOT | / | XRF, AR, MMI | MA |
| N | / | COMB | / | / | / | EA | / |
| Na | XRF, AR | MA | WD-XRF | WD-XRF | / | WD-XRF, AR | MA |
| Nb | TOT, AR, MMI | / | TOT | TOT | / | WD-XRF, AR, MMI | MA |
| Nd | TOT, AR, MMI | / | TOT | TOT | / | MMI | / |
| Ni | TOT, AR, MMI | / | TOT, AR | TOT, AR | TOT | WD-XRF, AR, MMI | MA |
| P | MMI, XRF | XRF | WD-XRF | WD-XRF |  | WD-XRF, AR, MMI | MA |
| Pb | TOT, AR, MMI | / | TOT, AR | TOT, AR | TOT | WD-XRF, AR, MMI | MA |
| Pd | FA, MMI | / | / | / | / | AR*, MMI* | / |
| Pr | TOT, AR, MMI | / | TOT | TOT |  | MMI | / |
| Pt | FA, MMI* | / | / | / | / | AR*, MMI* | / |
| Rb | TOT, AR, MMI | XRF | WD-XRF | WD-XRF | TOT | WD-XRF, AR, MMI | MA |
| Re | AR | / | / | / | / | / | / |
| S | XRF | XRF | AR | AR | / | EA, WD-XRF, AR, MMI | MA |
| Sb | TOT, AR, MMI | / | TOT | TOT | / | WD-XRF, AR, MMI | MA |
| Sc | TOT, AR, MMI | / | TOT | / | / | WD-XRF, AR, MMI | MA |
| Se | AR, MMI | / | / | / | / | AR, MMI | AAS |
| Si | XRF | XRF | WD-XRF | WD-XRF | / | WD-XRF | / |
| Sm | TOT, AR, MMI | / | TOT | TOT | / | MMI | / |
| Sn | TOT, AR, MMI | / | WD-XRF | WD-XRF | / | WD-XRF, AR, MMI | MA |
| Sr | TOT, AR, MMI | MA | WD-XRF | WD-XRF | TOT | WD-XRF, AR, MMI | MA |
| Ta | TOT, AR*, MMI | MA | TOT | TOT | / | WD-XRF, AR, MMI | / |
| Tb | TOT, AR, MMI | / | TOT | TOT | / | MMI | / |
| Te | AR, MMI* | / | TOT | TOT | / | AR, MMI | MA |
| Th | TOT, AR, MMI | / | TOT | WD-XRF | / | WD-XRF, AR, MMI | MA |
| Ti | MMI, XRF | XRF | WD-XRF | WD-XRF | / | WD-XRF, AR, MMI | MA |
| Tl | AR, MMI | MA | TOT | TOT | / | AR, MMI | MA |
| Tm | AR | / | / | / | / | / | / |
| U | TOT, AR, MMI | / | TOT | WD-XRF | / | WD-XRF, AR, MMI | MA |
| V | TOT, AR, MMI | / | TOT, AR | WD-XRF, AR | / | WD-XRF, AR, MMI | MA |
| W | TOT, AR, MMI | MA | WD-XRF | TOT | / | WD-XRF, AR, MMI | MA |
| Y | TOT, AR, MMI | / | WD-XRF | TOT | / | WD-XRF, AR, MMI | MA |
| Yb | TOT, AR, MMI | / | TOT | TOT | / | MMI | / |
| Zn | TOT, AR, MMI | / | WD-XRF, AR | WD-XRF, AR | TOT | WD-XRF, AR, MMI | MA |
| Zr | TOT, AR, MMI | XRF | WD-XRF | WD-XRF | / | WD-XRF, AR, MMI | / |
| Bulk Properties |  |  |  |  |  |  |  |
| CEC | / | / | / | / | / | STU | / |
| Clay content | LPSA | / | / | / | / | MIRS | / |
| EC | POT | / | / | / | / | / | / |
| LOI | CAL | GRAV, CAL | / | / | / | COMB | / |
| Mag Sus | / | / | / | / | / | SUS | / |
| Mineralogy | VNIRS | / | / | / | / | MIRS | XRD |
| pH | Field, POT | / | / | / | / | POT | / |
| Sand content | LPSA | / | / | / | / | MIRS | / |
| Silt content | LPSA | / | / | / | / | MIRS | / |
| TC | / | / | / | / | / | / | EA |
| TIC | / | / | / | / | / | / | XRD |
| TOC | / | / | LECO | IR |  | EA | CAL |
| Grain size 0.002 | / | / | LPSA | / | / | / | / |
| Grain size 0.06 | / | / | LPSA | / | / | / | / |
| Grain size D50% | / | / | LPSA | / | / | / | / |
| GrainsSortIndex | / | / | CAL | / | / | / | / |

1 Caritat et al. (2018); 2 Salminen et al. (2005) and Sandström et al. (2005); 3 Mixed sampling of outlet (floodplain) sediment and soil; 4 Residual soil; 5 Stream and Floodplain sediment (kept separate); 6 Agricultural and Grazing land soil (kept separate); 7 Soil. *All values below detection limit.

Notation:  AAS = Atomic Absorption Spectrometry; AFC = Alkaline Fusion Spectrophotometry; AFS = Atomic Fluorescence Spectrometry; AR = Aqua regia; CAL = Calculated; COMB = Combustion; DCA = Direct Current Arc; Emission Spectrometry; EA = Elemental Analyser; FA = Fire Assay; Field = Field pH test kit; GRAV = Gravimetry; IR = Non-dispersive infrared carbon analyser; ISE = Ion Specific Electrode; LPSA = Laser Particle Size Analyser; MA = Multi-Acid digestion (it is a four-acid digestion that includes HF); MIRS = Mid InfraRed Spectroscopy; MMI = Mobile Metal Ion®; POL = Polarography; POT = Potentiometry; STU = Silver-Thiourea extraction; SUS = Susceptibility meter; TIT = Titrimetry; TOT = Total digestion (fusion then digestion by HF + HNO3 after fusion); VNIRS = Visible Near InfraRed Spectroscopy; XRD = X-Ray Diffraction; XRF = X-Ray Fluorescence Spectrometry; ED-XRF = Energy Dispersive X-ray Spectrometry; WD-XRF = Wavelengh Dispersive X-ray Spectrometry; Grain size 0.002 = <0.002 mm (clay particle size); Grain size 0.06 = <0.06 to 0.002 mm (silt particle size); Grain size D50% = particle size when the cumulative percentage reaches 50%.

The analytical requirements to arrive at a comprehensive and satisfactory Global Geochemical Reference Network database need to be considered under two headings:

- Requirements for the Global Geochemical Reference Network, described in Chapter 3, and

- Additional requirements for national and international organisations engaged in the production of regional geochemical survey data.

National and regional geochemical surveys must be tied into the Global Geochemical Reference Network so that geochemical map data at any scale can be followed continuously across borders, from one jurisdiction to another (refer to Chapter 8 in this Manual). To achieve this, the Global Geochemical Reference Network and regional geochemical surveys must be tightly linked, methodologically and spatially, through:

- Shared secondary reference materials (SRMs).

- Common sampling points and sample media, and

- Matching analytical requirements.

There is a potential difficulty in that a small number of countries have already conducted a substantial amount of geochemical mapping to their own unique national or regional specifications. In some respects, these may differ from the recommendations contained in this Manual, which have been prepared to supply a common global perspective, and a standard for future geochemical surveys. Where the differences apply principally to analytical requirements, and original sample material has been retained, then it should be possible to reconcile differences by reanalysis. If original sample material has not been kept, or there are major differences in the types of sample material collected, a percentage of the original sample sites should be resampled according to these internationally standardised methods. There are serious difficulties in trying to reconcile data based on significantly different sample media (see Chapters 3 and 8 in this Manual). Actions must be determined on a case-by-case basis and cannot be further considered here. This work should be carried out by the project’s Laboratory, Quality Control and Data Management Committees.

## 6.2. Global Geochemical Reference Network analytical requirements

Samples which are to form part of a Global Geochemical Reference Network justify the greatest care and attention in all respects at all times. Data must be comparable. The analytical requirements for Global Geochemical Reference Network samples are governed by the following requirements, bearing in mind that some are not attainable in the short term for all elements, depending upon the sample media:

  - All elements in the periodic table, except H, O, Tc, Po, At, Fr, Ac, Pa and inert gases other than Rn, plus total loss-on-ignition should be determined (see Fig. 6.1).

  - The total amount of each element present is the most fundamental (and reproducible) quantity in any sample. Direct measurement techniques, e.g., X-ray fluorescence (XRF) or Instrumental Neutron Activation Analysis (INAA), or total extraction procedures using strong acid digestions and ICP-MS finish should, therefore, be employed as a priority. Accuracy is very important; the use of standard reference materials is discussed below (see Section §6.2.1).

  - Analytical methods that have detection limits significantly below the anticipated abundance of each element in the various specified sample media should be employed. Methods with insufficient sensitivity to produce reportable values for any given element in at least 80% of the samples cannot be considered satisfactory. It should be noted that the average abundance of a few elements (e.g., Ag, Au, Bi, Hg) in various media is not well-established.

  - High analytical precision is essential, and preferably significantly better than natural geochemical variation. The use of methods providing low detection limits gives the best assurance of good precision for trace elements. For all elements, however, it is the combined errors and uncertainties of site, sampling, preparation, and analysis that affect the significance of the data.

  - The analytical accuracy must also be good, preferably better than that achieved in most national regional geochemical surveys.

  - Non-destructive methods of analysis should be used where the above requirements can be met, to allow the use of large sample weights and reduce sample wastage, and

  - The amounts of the principal natural and anthropogenic gamma-emitting isotopes in the rock, residual soil, overbank and floodplain sediment samples should be determined, as a component of the geochemical baseline data.

It should be noted that samples from the Global Geochemical Reference Network sites could have the supplementary function of being used to study and monitor the dispersion of degradation products of organic contaminants in the environment, e.g., from pesticides and herbicides, solvents and plastic containers.

It is recommended that an evaluation of the best methods for the specific analytical requirements relating to the Global Geochemical Reference Network, should be determined by the Geoanalytical Committee of the Association of Applied Geochemists in collaboration with the project’s Laboratory Committee. From this evaluation, specific laboratories can be invited to participate in the analysis of samples collected for the Global Geochemical Reference Network.  Analytical schemes used in the pan-European continental-scale projects are described in Section §6.3. Recognising that the requirements, concerning the desired detection limits for some elements in some sample media, are beyond the capabilities of existing analytical procedures, every encouragement should be given to the appropriate authorities to sponsor the requisite research and development, which will enable the requirements to be met in the not-too-distant future.

Figure 6.1. Recommended list of elements for the Global Geochemical Reference Network project.

### 6.2.1. Use of standard reference materials

Analytical bias, both interlaboratory and batch-to-batch within a laboratory, must be reduced to a minimum to generate globally comparable data. This can be achieved by selecting the best analytical methods and using standard procedures in conjunction with routine monitoring of data quality. Primary (PRMs) and secondary (SRMs) reference materials must be prepared and used in these procedures (refer to Chapters 5, 7 & 8 in this Manual). The Canadian ‘STSD’ and Chinese ‘GSD’ (Xie et al., 2007) series of PRMs are internationally recognised stream sediment samples, well documented with recommended values of many elements. However, the quantities are not enough to be used in the Global Geochemical Reference Network project, and SRMs of each recommended sample type (rock, residual soil, stream sediment and overbank/floodplain sediment), with different element concentrations, must be prepared in large quantities to be used not only for the global project but also in national regional geochemical surveys (see Chapter 5).

Primary reference materials (PRMs) can be used for the selection of optimum analytical methods prior to the initiation of a geochemical mapping project. The preferred analytical method will be that which yields analytical data closest to the recommended values.

A set of PRMs should be inserted in each batch of 100 samples during routine multi-element analysis (approximately 4% of the total), i.e., usually at a rate of four PRMS at the start and end of each analytical batch (refer to Section §8.5 of Chapter 8 in this Manual). The analytical bias between different national geochemical surveys can be determined in this way and if necessary, the results from PRMs in each batch can be utilised for levelling data to an international standard reference level.

The number of SRMs inserted into a batch of 100 samples should be a minimum of two if the active project is collecting 1000s of samples (refer to Section §8.5 of Chapter 8).  If the project is collecting 100s of samples then this should be increased to four. If the number of samples collected is small (<100) then 1 in 20 samples and 1 in 10 samples should be an SRM. In Chapter 4, the quality control scheme requires the insertion of four SRMs in each batch of hundred samples of the Global Geochemical Reference Network project (see Table 4.1). It is stressed that the number of different SRMs submitted in an entire analytical batch of samples must be sufficient to produce satisfactory regression analysis plots (see Chapter 8 in this Manual).

The standard deviation of the determined values with the recommended values of PRMs and SRMs can be used to monitor the between-batch, between-method, between-map or between-laboratory bias. If the PRM and SRM data, obtained during any part of the work, are outside tolerance limits, then determinations should be repeated (Wisconsin Department of Natural Resources, 1996; EA, 1999, 2003a, 2003b, 2013; Eurachem, 2002; Ellison et al., 2003; 2019; Forbes, 2012; Magnusson and Örnemark, 2014; ISO, 2020a;).

### 6.2.2. Proposed arrangements for reference sample analysis

Because of the importance of establishing an internally consistent Global Geochemical Reference Network data set for each sample medium, it is strongly recommended to minimise or eliminate the problems that can arise in the inter-laboratory comparability of data, by engaging a small number of laboratories on this task, each selected for its demonstrated and recognised expertise in a particular method of analysis. Interlaboratory collaboration and cross-checks, using the same reference materials, should be installed.

If possible, a single laboratory should be responsible for one type of analysis performed on one sample type, collected worldwide, as has been done in the FOREGS Geochemical Atlas of Europe project (Salminen et al., 2005; Sandström et al., 2005). It is also desirable that where possible some elements should be determined by more than one analytical method to cross-check results.

It is discussed and recommended in Chapter 2 of this Manual that samples should be collected worldwide from approximately 7356 Global Terrestrial Network (GTN) grid cells of 160x160 km. From each GTN grid cell, samples shall be collected from either 5 or 8 random sites, following the principles outlined in Chapters 2 and 3. The media sampled will be as follows, and they are all mandatory:

- Rock

- Residual soil (top and C horizon)

- Humus

- Stream or lake sediment

- Stream water

- Overbank sediment (top and bottom)

- Floodplain sediment (top and bottom).

In an ideal landscape, from each quadrant of a 160x160 km grid cell seven sample types will be collected, i.e., five solid sample types, plus humus (if available), and stream water (if available). In total, between 25,000 and 36,000 samples of each type will be collected. Because humus and stream water samples will not be obtained in arid areas, there will be smaller numbers of these media. Lake sediments will be collected from less than 8% of the world’s surface. In order to apply ANOVA statistical procedures, it has been recommended that samples from each site should be analysed separately (see Chapter 7), because the cost of sample collection, particularly from remote areas, is substantial and the value of the additional information provided is considerable.

It is, therefore, recommended that each sample type should be sent to a designated receiving laboratory, which would coordinate all work on that particular sample type. Logically this laboratory could be responsible for applying one analytical technique, for example, XRF, INAA, ICP-AES or ICP-MS, on this particular sample type or possibly on all sample types. This laboratory would be responsible for distributing splits of the samples to other participating laboratories responsible for other analytical techniques. In this way, laboratories accepting responsibility for certain methods could be responsible, ultimately, for handling up to 72,000 samples of a given sample medium. Allowing for replication of analyses, the total number of analyses could be of the order of 5% greater than the number of sample sites for that particular sample medium. This volume of work is not considered excessive in terms of the number of samples commonly processed by production-orientated laboratories, given that samples will be received over several years. Even under ideal circumstances, with high priority and adequate funding, it is estimated that sample collection would take at least five years, provided all countries agree to participate in the Global Geochemical Reference Network project.

### 6.2.3. National and international geoanalytical requirements

The production of geochemical maps at scales commonly ranging from 1:1 million to 1:50,000 is normally the responsibility of national organisations. In addition to providing a global overview, the prime reason for the establishment of the Global Geochemical Reference Network is to make it possible for researchers to compare and relate geochemical phenomena anywhere in the world.

It is proposed that a split of analysed reference samples from each grid cell in the Global Geochemical Reference Network should be retained in the country or region of origin, for ongoing reference purposes and quality control of subsequent geochemical work in that country.  These sample splits will serve as secondary reference materials. As analytical data for these reference samples become available, a copy will be forwarded to the responsible national earth science institutes in the countries of origin. The existence of reference data and materials relevant to every region will make it possible for neighbouring countries to compare and connect the results of their detailed geochemical survey work and contribute to part of a single international database.

Recommended field sampling methods to encourage greater commonality for regional geochemical surveys are given in Chapters 2 and 3 in this Manual. Concerning the analytical requirements for these surveys, it is assumed that many countries will wish to develop and use their in-house analytical capabilities. Those not in a position to do so should, as an alternative, provide specifications for the work to be carried out by others, but these specifications should be compatible with recommended international procedures and employ the locally available reference materials.

The existence of international recommendations regarding the selection of sample media and analytical methods is not intended to suggest that organisations should not choose what they consider to be the optimum sample type, sample spacing or analytical method for their own particular purposes. However, it is strongly recommended that whenever other sample media or analytical methods are chosen by a country, these should be linked to the Global Geochemical Reference Network, by sampling at the same sites as are used for the Global Geochemical Reference Network and using the global project’s secondary reference materials to establish a known relationship between their national analytical procedures and those recommended for international use.

It is recommended above that for the Global Geochemical Reference Network is desirable, ultimately, to obtain data for almost the entire periodic table. It is unlikely that many countries will see a need to make detailed maps for the complete list. Nevertheless, given the heterogeneous collection of data in existing data sets, it is highly desirable that a common range of elements of known economic and environmental importance should be determined in every country by internationally agreed standard methods. If some of the necessary facilities are lacking, rather than accept inferior data, consideration should be given to having some of the work undertaken in an internationally recognised laboratory. The recommended elements to be determined are grouped into ‘analytical packages’ (Fig. 6.1):

- List 1 elements, for economic and environmental purposes, including a group of major elements and a group of minor, trace and ultra-trace elements (total 55 determinands) that should be determined in any new geochemical mapping project, and

- List 2 elements that could usefully be added depending upon the funding of the project and the availability of laboratory facilities (total of 24 determinands).

Analytical precision will depend upon the method used and the element determined. Generally speaking, the relative standard deviation (RSD) for the major elements should be less than 3%, the RSD for minor and trace elements should be less than 10%, and the RSD for certain trace and ultra-trace elements should be less than 20%.

### 6.2.4. Need for independent quality control

Communication between laboratory and field staff is of paramount importance. Users of analytical data must always check the validity of the results reported and those generating the analytical data should be aware of the purposes for which the data are being utilised. Blind confidence should not be placed upon any laboratory’s quality assurance and control programme. An external quality assurance and control procedure must be installed by the Quality Control Committee of the Global Geochemical Reference Network project.

Laboratory quality assurance and control programmes may not always suit the types of sample matrices submitted or are rigorous enough to satisfy requirements close to detection limits. Normally, laboratories evaluate a new method by its application to well-characterised PRMs, available commercially from a number of organisations such as CANMET Canada; NIST (formerly the U.S. National Bureau of Standards), the United States Geological Survey Geochemical Reference Materials, International Association of Geoanalysts, LGC Geological materials, Institute of Geophysical and Geochemical Exploration, China (Xie et al., 2007). They then standardise their own internal controls (SRMs) against these for frequent insertion in sample batches. For regional geochemical mapping, and especially the Global Geochemical Reference Network project, the organisations responsible should develop and standardise their own bulk control samples (SRMs), the matrices of which are typical of those found in the region under study. Replicate data for these control samples will allow computation of precision and accuracy of the methods used, and also to level the results of disparate data sets (refer to Chapter 8 in this Manual). To assess the relative magnitude of analytical variability (precision) versus the spatial geochemical variability and bias introduced during sampling, replicate analyses of the survey samples should be carried out and Analyses of Variance determined (see Chapter 7 in this Manual).

## 6.3. Analytical schemes used in multinational continental-scale projects

The Geochemical Atlas of Europe of the Forum of European Geological Surveys (FOREGS presently EuroGeoSurveys), known as the FOREGS Atlas is widely regarded as a pilot for the Global Geochemical Reference Network project as it is the first multi-national and multi-media project in which 26 countries participated (Salminen et al., 2005; De Vos, Tarvainen et al., 2006), and followed the recommendations of Darnley et al. (1995) for the development of an internally consistent “global geochemical database for environmental and resource management”. Such a continental-scale project demands special requirements for chemical analysis, which have already been mentioned in Section §6.2. Nevertheless, it is important to stress that the accuracy of the European data set was essential since one of the objectives, according to Darnley et al. (1995), is for the results to be used to ‘normalise’ or to ‘level’ other geochemical data at national and international scales (refer to Chapter 2 in this Manual and its two Annexes). The methods used in this project are also likely to become standard operating procedures (SOPs) for the Global Geochemical Reference Network project, and hence should be reproducible in various laboratories throughout the world. Because of the stringent requirements regarding sensitivity and elemental coverage, the most sophisticated instrumental methods have to be utilised. This may restrict some skilled but poorly equipped national laboratories from participating fully in the global project in the future.

The material of this Section was extracted from the FOREGS Analytical Committee’s report prepared by Sandström et al. (2005), and modified for the purposes of this Manual.

Total element concentrations are most relevant for the geochemical interpretation of data. For solid materials, this means that the silicate matrix either needs to be fully decomposed by mixed acid digestion before instrumental analysis, or a solid sampling technique such as X-ray fluorescence needs to be used. However, to address the needs of national and European level environmental authorities, information on leachable concentrations of the elements was also considered to be important. In environmental chemistry, a slightly unscientific and non-specific term, ‘near total’, is often used to describe the maximum concentration of an element that can be liberated from a material in its natural environment. Aqua regia digestion is commonly used for simulating this process in the laboratory. Unfortunately, almost every laboratory has its own standard operating procedure for aqua regia digestion. For this project, an aqua regia leach was adopted to determine the leachable concentrations of the elements in sediment and soil samples. Errors introduced by a diversity of leaching procedures were avoided by having only one laboratory perform the extractions and analyses. The method used was validated against standard methods according to normal validation procedures in accredited laboratories.

### 6.3.1. Participating laboratories and analytical techniques

Nine laboratories were involved in the FOREGS Atlas project (Table 6.2). All solid samples collected by the 26 participating countries were shipped to the Slovakian State Geological Institute of Dionyz Stur (SGUDS), which was in overall charge of sample preparation, homogenisation, sub-sampling, and sample distribution to the analytical laboratories and archiving of the excess sample material (see Chapter 4 ‘Sample Preparation and Storage’). To ensure data homogeneity, and to avoid any bias between laboratories or analytical methods, each laboratory was nominated to take responsibility for carrying out analysis by a particular analytical technique or techniques on all samples of a certain type. Some determinands were measured by more than one technique, e.g., a range of trace elements in samples of soil and sediment were determined both by XRF and ICP-MS after mixed acid digestion. The laboratories and actual methods used were selected through careful consideration of a few quality issues and are listed in Table 6.3.

Table 6.2. European Geological Survey laboratories participating in the FOREGS Geochemical Atlas of Europe project.

| Laboratory | Acronym |
| --- | --- |
| Bundesanstalt für Geowissenschaften und Rohstoffe, Hannover, Germany | BGR |
| British Geological Survey, Keyworth, Nottingham, United Kingdom | BGS |
| Bureau de Recherches Géologiques et Minières, Orleans-Cedex, France | BRGM |
| State Geological Institute of Dionyz Stur, Bratislava, Slovak Republic | SGUDS |
| Geological Survey of Finland, Espoo, Finland | GTK |
| Mining and Geological Survey of Hungary, Budapest, Hungary | MBFSZ |
| Geological Survey of The Netherlands, Utrecht, The Netherlands | TNO |
| Geological Survey of Norway, Trondheim, Norway | NGU |
| Polish Geological Institute, Warszawa, Poland | PGI |

Table 6.3. Summary of analytical methods used for each sampling medium in the FOREGS Geochemical Atlas of Europe project.

| Topsoil and Subsoil samples |  | Stream and Floodplain Sediment samples |  |  |
| --- | --- | --- | --- | --- |
| Laboratory | Method | Laboratory |  | Method |
| SGUDS | Sample Preparation | SGUDS |  | Sample Preparation |
| TNO | Granulometry/TOC | SGUDS |  | TOC |
| BGR | WD-XRFS | BGS |  | XRFS |
| GTK | ICP-MS | BRGM |  | ICP-MS |
| PGI | AR/ICP-AES | GTK/PGI |  | AR/ICP-AES |
| MBFSZ | Hg analyser | MBFSZ |  | Hg analyser |
| Stream Water samples |  | Humus samples |  |  |
| Laboratory | Method | Laboratory | Method |  |
| BGR | ICP-QMS / ICP-AES | SGUDS | Sample preparation |  |
| BGS | IC / DOC analyser | TNO | ICP-MS |  |
| BGS | DOC analyser | MBFSZ | Hg analyser |  |
| NGU | Hg analyser |  |  |  |

### 6.3.2. Sample preparation

Refer to Chapter 4 in this Manual.

### 6.3.3. Analytical methods used in the European Global Terrestrial Network

#### 6.3.3.1. Analysis of stream water samples

##### 6.3.3.1.1. Determination of anions

Anions were determined on stream water samples that were unfiltered and unpreserved but had been stored after sampling at between 1 and 8°C prior to analysis. Stream water samples were analysed for major and trace anions, i.e., fluoride, chloride, nitrite, bromide, nitrate, orthophosphate and sulphate, by ion chromatography (IC). The determinations were made on a Dionex 4000i Ion Chromatograph using an AS14 analytical column with both conductivity and UV/visible absorbance detection. The method is accredited by the United Kingdom Accreditation Service (UKAS) according to the requirements of BS EN ISO 17025 (ISO, 2017).

The stream water samples were filtered to <0.45 µm before analysis. After chromatographic separation, all the anions were measured using the conductivity detector. Nitrite, bromide and nitrate were also determined using the UV/visible absorption detector, and data for these anions were taken from this detector because of its better sensitivity.

For all anions, the accuracy of the method was better than ±10%, the bias was within ±3% and the precision at the 95% confidence interval was better than 5% at concentrations an order of magnitude above the quantification limit. Quantification limits (QL) for all anions are given in Table 6.4.

##### 6.3.3.1.2. Determination of dissolved organic carbon

Dissolved organic carbon (DOC) was determined on stream water samples that were filtered to <0.45 µm on collection but were otherwise unpreserved. Stream water samples were analysed using a Shimadzu TOC 5000 analyser with associated ASI 5000 auto-sampler. The method is accredited by the United Kingdom Accreditation Service according to the requirements of BS EN ISO 17025 (ISO, 2017).

Stream water samples were acidified and sparged prior to analysis to remove inorganic carbon. DOC was then determined by thermal combustion of the sample and the evolved CO2 was measured by non-dispersive infra-red (NDIR) detection.

The accuracy of the method was better than ±5%, the bias was within ±2%, and the precision at the 95% confidence interval was better than 6% at concentrations an order of magnitude above the quantification limit at 0.5 mg/l.

##### 6.3.3.1.3. Determination of major cations, metals and other inorganic elements

Stream water samples were filtered to <0.45 µm, acidified to 1% v/v with nitric acid and stored at a temperature of <8°C. The samples were analysed by both inductively coupled plasma quadrupole mass spectrometry (ICP-QMS) and inductively coupled plasma atomic emission spectrometry (ICP-AES), using Perkin Elmer Sciex ELAN 5000A and Spectro Flame M instruments, respectively, in accordance with the German norms DIN 38406-29 (ICP-MS) and DIN 38406-22 (ICP-AES) (DIN, 1988, 1996).

The primary chemicals used to prepare the calibration and quality control standards and reagents were of analytical reagent grade. Multielement standard solutions for calibration were prepared from Claritas SPEX/Certiprep stock solutions. International certified reference materials (NIST 1640, NIST 1643d and SLSR-4) were included in every batch of 20 samples.  Indium was used as an internal standard.

The accuracy of the methods for all determinands was better than ±10%, the bias was within ±3%, and the repeatability at the 95% confidence interval was better than 5% at concentrations an order of magnitude above the quantification limit. Quantification limits for all cations and trace elements are given in Table 6.4.

##### 6.3.3.1.4. Determination of mercury

Stream water samples were filtered to less than 0.45 µm and preserved in the field using potassium dichromate dissolved in HNO3. The samples were analysed for mercury by a cold vapour atomic absorption (CV-AAS) technique using a CETAC M-6000A Hg Analyser. The method is accredited according to NS-EN ISO/IEC 17025 (ISO, 2017).

The method is based on the reduction of Hg ions in solution to atomic Hg vapour using stannous chloride (SnCl2) as a reducing agent. The Hg/SnCl2 emulsion is introduced into the top of a gas-liquid separator and is detected by CV-AAS.

The accuracy of the method was estimated to be ±1 to 2%, without any systematic bias.  Within batch precision at the 95% confidence interval was estimated to be about 1%, and long-term precision better than 8%. The quantification limit was 0.01 µg/l.

Table 6.4. Summary of all elements and other determinands in the FOREGS Geochemical Atlas of Europe (after Sandström et al., 2005) for each sampling medium, including the analytical method and laboratory responsible for analysis with the limit of quantification, and the Upper Continental Crust abundance for comparison. Units are denoted by superscripts: amg/kg;  b%;  cµg/l;  dmg/l;  eμg/g;  fmS/m;  gmm;  hμm

| Element | Soil samples |  | Sediment samples |  | Humus | Stream water samples | Upper Continental Crust abundance (mg/kg) | Source |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Total | Aqua regia | Total | Aqua regia |  |  |  |  |
| Ag | 0.01a | / | / | / | / | 0.002c | 0.053 | Rudnick and Gao (2003) |
| Al / Al2O3 | 0.05b | / | 0.1b | / | / | 0.1c | 81,505 | Rudnick and Gao (2003) |
| As | 0.2a | 5a | 1a | 5a | / | 0.01c | 5.7 | Hu and Gao (2008) |
| B | / | / | / | / | / | 0.01c | 47 | Hu and Gao (2008) |
| Ba | 5a | 1a | 3a | 1a | 1a | 0.005c | 628 | Rudnick and Gao (2003) |
| Be | 2a | / | 0.02a | / | / | 0.005c | 1.9 | Hu and Gao (2008) |
| Bi | 0.5a | / | / | / | / | 0.002c | 0.23 | Hu and Gao (2008) |
| Ca / CaO | 0.01b | / | 0.05b | / | / | 0.001d | 25,657 | Rudnick and Gao (2003) |
| Cd | 0.01a | / | 0.02a | / | 0.1a | 0.002c | 0.06 | Hu and Gao (2008) |
| Ce | 0.15a | / | 0.02a | / | / | 0.002c | 63 | Rudnick and Gao (2003) |
| Co | 3a | 1a | 2a | 1a | 0.1a | 0.005c | 15 | Hu and Gao (2008) |
| Cr | 3a | 1a | 3a | 1a | / | 0.01c | 73 | Hu and Gao (2008) |
| Cs | 0.5a | / | 4a | / | / | 0.002c | 4.9 | Hu and Gao (2008) |
| Cu | 0.01a | 1a | 1a | 1a | 0.3a | 0.005c | 27 | Hu and Gao (2008) |
| Dy | 0.1a | / | 0.02a | / | / | 0.002c | 3.9 | Rudnick and Gao (2003) |
| Er | 0.1a | / | 0.02a | / | / | 0.002c | 2.3 | Hu and Gao (2008) |
| Eu | 0.05a | / | 0.02a | / | / | 0.002c | 1 | Rudnick and Gao (2003) |
| Fe / Fe2O3 | 0.01b | 2a | 0.01b | 2a | / | 1c | 39,176 | Rudnick and Gao (2003) |
| Ga | 0.2a | / | 1a | / | 0.1a | 0.002c | 18.6 | Hu and Gao (2008) |
| Gd | 0.1a | / | 0.02a | / | / | 0.002c | 4 | Rudnick and Gao (2003) |
| Ge | / | / | / | / | / | 0.005c | 1.3 | Hu and Gao (2008) |
| Hf | 0.2a | / | 0.05a | / | / | 0.002c | 5.3 | Rudnick and Gao (2003) |
| Hg | 0.0001e | / | 0.0001e | / | 0.0001e | / | 0.05 | Rudnick and Gao (2003) |
| Ho | 0.02a | / | 0.02a | / | / | 0.002c | 0.83 | Rudnick and Gao (2003) |
| I | 2a | / | / | / | / | 0.01c | 1.4 | Rudnick and Gao (2003) |
| In | 0.01a | / | / | / | / | 0.002c | 0.066 | Hu and Gao (2008) |
| La | 0.01a |  | 0.02a |  | 0.1a | 0.002c | 31 | Rudnick and Gao (2003) |
| Li | / | / | 0.05a | / | / | 0.005c | 41 | Hu and Gao (2008) |
| Lu | 0.02a | / | 0.02a | / | / | 0.002c | 0.36 | Hu and Gao (2008) |
| Mg / MgO | 0.01b | / | 0.1b | / | / | 0.001d | 14,955 | Rudnick and Gao (2003) |
| Mn / MnO | 0.001b | 10a | 0.01b | 10a | / | 0.05c | 774 | Rudnick and Gao (2003) |
| Mo | 0.1a | / | 0.05a | / | / | 0.002c | 0.6 | Hu and Gao (2008) |
| Na / Na2O | 0.01b | / | 0.2b | / | / | 0.005d | 24,259 | Rudnick and Gao (2003) |
| Nb | 0.1a | / | 1a | / | / | 0.002c | 11.6 | Hu and Gao (2008) |
| Nd | 0.15a | / | 0.02a | / | / | 0.005c | 27 | Rudnick and Gao (2003) |
| Ni | 2a | 2a | 1a | 2a | 0.3a | 0.005c | 34 | Hu and Gao (2008) |
| P2O5 | 0.001b | / | 0.01b | / | / | / | 655 | Rudnick and Gao (2003) |
| Pb | 3a | 3a | 1a | 3a | 0.5a | 0.005c | 17 | Rudnick and Gao (2003) |
| Pr | 0.1a | / | 0.02a | / | / | 0.002c | 7.1 | Rudnick and Gao (2003) |
| Rb | 2a | / | 1a | / | 0.2a | 0.002c | 94 | Hu and Gao (2008) |
| S | / | 50a | / | 50a | / | / | 62 | Rudnick and Gao (2003) |
| Sb | 0.02a | / | 0.02a | / | / | 0.002c | 0.75 | Hu and Gao (2008) |
| Sc | 0.5a | / | / | / | / | / | 14 | Hu and Gao (2008) |
| Se | / | / | / | / | / | 0.01c | 0.09 | Rudnick and Gao (2003) |
| SiO2 | 0.1b | / | 0.1b | / | / | 0.01d | 311,405 | Rudnick and Gao (2003) |
| Sm | 0.1a | / | 0.02a | / | / | 0.002c | 4.7 | Rudnick and Gao (2003) |
| Sn | 2a | / | 1a | / | / | / | 2.2 | Hu and Gao (2008) |
| Sr | 2a | / | 1a | / | 1a | 0.001d | 320 | Rudnick and Gao (2003) |
| Ta | 0.05a | / | 0.05a | / | / | 0.002c | 0.92 | Hu and Gao (2008) |
| Tb | 0.02a | / | 0.02a | / | / | 0.002c | 0.7 | Rudnick and Gao (2003) |
| Te | 0.02a | / | / | / | / | 0.005c | 0.027 | Hu and Gao (2008) |
| Th | 0.1a | / | 1a | / | / | 0.002c | 10.5 | Rudnick and Gao (2003) |
| Ti / TiO2 | 0.001b | / | 0.005b | / | / | 0.01c | 3836 | Rudnick and Gao (2003) |
| Tl | 0.01a | / | 0.02a | / | / | 0.002c | 0.55 | Hu and Gao (2008) |
| Tm | 0.02a | / | 0.02a | / | / | 0.002c | 0.37 | Hu and Gao (2008) |
| U | 0.1a | / | 1a | / | / | 0.002c | 2.6 | Hu and Gao (2008) |
| V | 0.5a | 1a | 2a | 1a | / | 0.05c | 106 | Hu and Gao (2008) |
| W | 5a | / | 0.05a | / | / | 0.002c | 1.4 | Hu and Gao (2008) |
| Y | 3a | / | 0.02a | / | / | 0.002c | 21 | Rudnick and Gao (2003) |
| Yb | 0.05a | / | 0.02a | / | / | 0.002c | 2.34 | Hu and Gao (2008) |
| Zn | 3a | 1a | 1a | 1a | 1a | 0.01c | 75 | Hu and Gao (2008) |
| Zr | 3a | / | 1a | / | / | 0.002c | 193 | Rudnick and Gao (2003) |
| EC | / | / | / | / | / | f |  |  |
| HCO3- | / | / | / | / | / | d |  |  |
| Br- | / | / | / | / | / | 0.01d |  |  |
| Cl- | / | / | / | / | / | 0.1d |  |  |
| F- | / | / | / | / | / | 0.05d |  |  |
| NO3- | / | / | / | / | / | 0.04d |  |  |
| SO42- | / | / | / | / | / | 0.3d |  |  |
| DOC | / | / | / | / | / | c |  |  |
| TOC | b | / | b | / | / | / |  |  |
| Grain size 0.002g | b | / | / | / | / | / |  |  |
| Grain size 0.06g | b | / | / | / | / | / |  |  |
| Grain size D50%h | b | / | / | / | / | / |  |  |
| GrainsSortIndex | unitless | / | / | / | / | / |  |  |
| Units: amg/kg; b%; cµg/l; dmg/l; eμg/g; fmS/m; gmm; hμm |  |  |  |  |  |  |  |  |
| Colours: | Refer to Table 6.2 for laboratory acronyms. |  |  |  |  |  |  |  |
| ICP-MS | GTK | BGR | BRGM | TNO |  |  |  |  |
| ICP-AES | PGI | BGR |  |  |  |  |  |  |
| XRF | BGS | BGR |  |  |  |  |  |  |
| IC | BGS |  |  |  |  |  |  |  |
| CV-AAS | MFGI |  |  |  |  |  |  |  |

##### 6.3.3.1.5. Selection of stream water data

Data for the major cations (Ca, Mg, Na and K) and some of the more abundant trace elements (Sr, Mn and Al), as well as Si and total P, were taken from the ICP-AES data set. All other trace element data were obtained from the ICP-MS data set. A range of trace elements was analysed by both ICP-AES and ICP-MS. The data were compared to ensure that there was good agreement between the two data sets. The ICP-MS data were generally selected for publication because of the better sensitivity of this instrumental method. All cation and trace element data determined by ICP techniques were considered to be of acceptable quality, except for Ag. Contamination, probably arising from the poor quality of the acid preservative used in a few countries, was the likely cause of the poor quality of Ag data.

Major anion data determined by IC, as well as the TOC data, were all considered to be of good quality. Because of the variable length of time between collection of the stream water samples and analysis in the laboratory, there is evidence that nitrite and phosphate, which are relatively unstable, may have been lost from solution. Data for nitrite and phosphate are not considered, therefore, to be acceptable for publication. The detection limits for bromide and phosphate were also not sufficiently sensitive for this study, and less than 10% of the samples contained concentrations of these anions above quantifiable levels.

Because of contamination, caused by inappropriate sample bottles being used in some countries, the mercury data were, unfortunately, not accepted for publication. Hence, it is very important for the Project Manager of the Global Geochemical Reference Network project to ensure that the centrally purchased stream water sample bottles are certified trace element free, and to be distributed to all participating countries.

Data for major cations determined by ICP-AES and anions determined by IC were checked for accuracy by calculating an ionic charge balance and also by checking the ratio of conductivity to total dissolved solids (TDS). Unfortunately, the major anion in most water samples was bicarbonate, which was determined by each country in the field. The quality of the bicarbonate data was highly variable, including errors in the analytical determination, errors in reporting units and some data were also not available. Nevertheless, a few rules and additional checks were established and applied to the stream water data set that enabled a thorough check on the overall quality of the analyses. This is another standardisation problem that should be addressed by all participating countries: the same field instruments and the same chemical reagents must be used, and most importantly all field sampling teams must be trained in their use.

Further details of the full range of elements determined are given in Table 6.4.

#### 6.3.3.2. Analysis of stream and floodplain sediment samples

##### 6.3.3.2.1. Determination of TOC

Determination of total carbon content in stream and floodplain sediment <0.063 mm samples was carried out by a pyrolysis technique with non-dispersive infrared detection using a Dohrmann-Rosemount DC-190 (USA) carbon analyser, with a 183-boat sampling module and a Fuji model 3300 direct non-dispersive infrared gas detector. The method has been validated against USEPA methods 415.1 and 9060 (USEPA, 1997, 2004) and ISO 10694 (ISO, 1995).

Well-mixed samples were treated with 50% HNO3, and organic compounds were decomposed by pyrolysis in the presence of oxygen at a temperature of 800°C.

The uncertainty of the method was estimated to be ±10% and was mainly limited by sample heterogeneity.

##### 6.3.3.2.2. Determination of mercury

Determination of total mercury was carried out using a cold vapour atomic absorption technique (CV-AAS), involving preconcentration on a gold amalgam before detection with an Advanced Mercury Analyser (AMA-254, ALTEC) instrument.

The analysis was performed directly on solid <0.063 mm samples without any sample preparation. Mercury is liberated from the sample during programmed temperature elevation to 850oC and amalgamated with gold to concentrate the determinand. Mercury vapour is then released from the amalgam by heating and detected using atomic absorption spectrometry.

The uncertainty of the analysis was typically better than ±10%, or better than 5% in the case of higher concentrations. The quantification limit of Hg was 0.0001 µg/g.

##### 6.3.3.2.3. Total multielement determination by XRF

A range of elements was determined by wavelength dispersive X-ray fluorescence spectrometry (WD-XRFS) and energy dispersive polarised X-ray fluorescence spectrometry (ED(P)XRFS).  The instruments used were Philips PW1480 and PW2400 WD-XRFs, with W and Rh anode X-ray tubes respectively, and a Spectro X-LAB 2000 ED-XRF with a Pd anode Χ-ray tube.

Samples (<0.063 mm) were prepared by mixing with a binder and pressed into powder pellets before analysis by both instrumental techniques. For the WD methods, secondary radiation is collimated onto a diffraction crystal and its intensity at selected peak and background positions in the X-ray spectrum is measured using a detector mounted onto a goniometer. The net intensity, corrected for background, spectral interference and instrument drift, is calibrated against known synthetic standards and certified reference materials (CRMs). For the ED(P)XRFS method, five different secondary/polarisation targets are used to give optimal coverage of 52 elements from Na to U; all elements are measured to improve the accuracy of the corrections on the analytes of interest. Rather than being diffracted, the whole of the emitted X-ray spectrum is detected simultaneously using a Si(Li) detector. The acquired spectrum is deconvoluted and then evaluated using a calibration prepared by the instrument manufacturer. In general, data for lighter elements are taken from the WD-XRF instrument, and data for the heavier elements from the ED(P)XRFS instrument.

In practice, concentrations of MgO, P2O5, K2O, CaO, TiO2, V, Cr, MnO, Cs, Ba, La and Ce were obtained by using the ED technique. The WD technique was used for all other elements.  Further details of the full range of elements determined are given in Table 6.4.

The accuracy of the method for all determinands was better than ±5%, and the repeatability at the 95% confidence interval was better than 5% at concentrations an order of magnitude above the limit of quantification. Quantification limits for all determinands are given in Table 6.4.

##### 6.3.3.2.4. Total multielement determination by ICP-MS

In addition to XRF analysis, a range of major and trace elements were also determined by quadrupole ICP-MS using a VG Elemental (UK) model PQ3 instrument equipped with a water- cooled sample introduction system, a PlasmaScreen torch, a high-efficiency interface device AutoRange Plus, a simultaneous detector system and PlasmaLab ICP-MS Software Suite.

Prior to analysis, the stream and floodplain sediment samples were decomposed by sintering with sodium peroxide (1 g of <0.063 mm sample and 3 g flux) at a temperature of 480°C. After cooling, the sinter was dissolved in 60 ml of deionised water and 20 ml of 50% nitric acid. Rhodium was used as an internal standard. Nowadays, indium is commonly used as an internal standard because of its lower cost.

At concentrations an order of magnitude above the quantification limit, the uncertainty of the method was between 5 and 10%, depending on the element. Quantification limits are given in Table 6.4.

##### 6.3.3.2.5. Acid leachable multielement determination by ICP-AES

In addition to the total element concentrations, the acid leachable portion of selected elements was analysed after a hot aqua regia leach by ICP-AES, using a J-Y 70 Plus Geoplasma ICP-AES instrument.

Stream and floodplain sediment samples were digested in aqua regia by weighing 1 g of <0.063 mm sample in a polyethylene tube and adding 6 ml HCl and 2 ml HNO3. The samples were left for 15 minutes at room temperature before heating in an aluminium block at 95°C for 60 minutes. After cooling, the samples were filtered and made up to 50 ml in a polyethylene flask.

Based on monitoring samples, the relative standard deviation of the measurements was typically between 3 and 7% for concentrations exceeding three times the quantification limit of the particular element. Quantification limits for all reported elements are given in Table 6.4.

##### 6.3.3.2.6. Selection of stream and floodplain sediment data

A wide range of total multielement data determined by both ICP-MS and XRFS methods were available. Data quality, including results of CRMs and ANOVA data for duplicate-replicate analyses, was assessed on an element-by-element basis. If only one of the two techniques were sufficiently sensitive to permit the determination of a particular element at the background levels found in the samples, then this technique was automatically selected. If both techniques were sufficiently sensitive, the data from the technique with the best precision, based on the ANOVA duplicate analyses, were selected. The actual method selected for reporting each element is given in Table 6.4.

The quality of the ANOVA data for the ICP-AES analysis of the aqua regia solutions was also assessed. For a few elements, the analytical variation exceeded that of the geochemical variation, and the data for these elements were not considered suitable for publication. Those elements with acceptable data quality are given in Table 6.4. Aqua regia data were also used for total S, for which no data were available by any other technique.

#### 6.3.3.3. Analysis of residual soil samples

##### 6.3.3.3.1. Determination of particle size distribution

Particle size (grain size) distribution between 0.1 and 2000 μm was measured by a Laser Particle Sizer (LPS) technique using a Malvern Mastersizer 2000 instrument (Malvern, UK) equipped with a Hydro 2000G sample dispersion accessory. The method is based on the correlation between the angles of light scattered from particles in a laser beam and the size distribution of these particles.

Latex spheres with unique known grain sizes traceable to NIST (Malvern Quality Audit Standard QAS2002) were used to calibrate the Laser Particle Sizer.

Samples were sieved to <2 mm, followed by the removal of organic matter and shells by oxidising the sample with 10 ml 30% H2O2 and heating to 70°C. Carbonates and iron oxides were removed by adding 10% HCl and heating. A suspension of the sample in Na4P2O7.10H2O was measured with a Laser Particle Sizer instrument.

The grain size distribution is expressed as the measured volume divided into 30 fractions, logarithmically spaced between 0.1 and 2000 µm. The software program is also able to calculate the different distribution classes (D (0.5), D (0.6), etc.).

##### 6.3.3.3.2. Determination of TOC

The total organic carbon (TOC) content of <0.063 mm soil samples was analysed using a LECO SC-DR144 instrument coupled to a Mettler AT 400 analytical balance.

The method is based on the total combustion of the samples after the removal of the carbonate phase by hydrochloric acid. After combustion in a 99.99 % oxygen flow at 1350°C, the resulting CO2 is passed through a water and halogen trap and detected with an infrared detector. Analytical grade CaCO3 was used to construct a four-point linear calibration.

The detection limit for total carbon was estimated to be 0.01 wt% and the uncertainty was estimated to be better than ±5%.

##### 6.3.3.3.3. Determination of mercury

Mercury in residual soil samples was determined by the method described in Section §6.3.3.2.2 of ‘Stream and floodplain sediments’.

##### 6.3.3.3.4. Total multielement determination by XRF

A range of elements was determined by wavelength dispersive X-ray fluorescence spectrometry (WD-XRFS) using Philips PW1480 and PW2400 WD-XRF instruments, with Cr and Rh anode X-ray tubes, respectively.

Residual soil samples were prepared by milling in an agate mortar to <40 µm (0.04 mm) particle size. Then 1000 mg per sample were mixed with 5.0 g lithium metaborate and 25 mg lithium bromide and fused at 1200°C for 20 minutes. Loss on ignition (LOI) was determined by heating to 1030°C for 10 minutes. For samples with a LOI greater than 20%, 2.5 g lithium metaborate and 2.415 g lithium tetraborate were used. Pt95-Au5 crucibles and a commercial automatic fluxer (Herzog 12/1500) were used for the fusion.

The calibration was carried out using 130 certified reference materials (CRMs) and corrections were made for matrix interferences. In addition to the routine laboratory quality control protocol, LUFA-B (a German residual soil reference sample) was analysed regularly throughout the batches of samples submitted for analysis.

##### 6.3.3.3.5. Total multielement determination by ICP-MS

A range of rare earth elements (REE) and other trace elements was determined on residual soil samples using a Perkin-Elmer Sciex Elan 5000 inductively coupled plasma mass spectrometer.

Two sample digestion methods were used depending on the elements to be analysed in order to optimise the elemental recovery and detection limits. In the main method, 0.2 g of the <0.063 mm soil sample was weighed into a Teflon dish and the organic material totally decomposed by evaporating the sample to dryness with 5 ml of 65% nitric acid, followed by the addition of 10 ml of 40% hydrofluoric and 4 ml of 70% perchloric acid and evaporated on a hot plate. The residue was dissolved in 20 ml of 8 mol/l nitric acid and 1 ml of 30% hydrogen peroxide before filtration. The filtrate was saved, and the filter paper was ashed in a platinum crucible. The residue was then fused with 0.2 g of lithium metaborate, and 0.02 g of sodium perborate followed by dissolution in 5 ml of 0.8 mol/l nitric acid. The solutions were combined and made up to 100 ml in 1.8 mol/l nitric acid. Measurement of these digests was carried out using a Perkin Elmer Sciex Elan 5000 inductively coupled plasma-mass spectrometer.

A simple cold dissolution method was used for selected elements: As, Cd, Cs, Cu, Mo, Ni, Pb and Sn. In this method, 0.1 g of the <0.063 mm soil sample was weighed into a plastic test tube and 1.5 ml of nitric acid and 0.5 ml of hydrofluoric acid were added. The test tube was closed with a stopper, agitated and allowed to stand at room temperature for 48 h. Deionised water (8 ml) was then added and the solution was allowed to stand for a further 48 h. The test tubes were then agitated and centrifuged, and the solution was filtered. The solution was diluted before analysis using a Perkin Elmer Sciex Elan 6000 inductively coupled plasma-mass spectrometer.

A reagent blank and two CRMs (Slovakian Residual Soil 982 and River Clay 981) were included as monitoring samples in every batch of 40 samples. In addition, CRMs were included within each batch of 80 samples for the main digestion method (syenite SY-2) and within each batch of 40 samples for the cold dissolution method (San Joaquin Soil NIST 2701).

At concentrations an order of magnitude above the quantification limit, the uncertainty of the method was between 1 and 10%, depending on the element. Quantification limits are given in Table 6.4.

##### 6.3.3.3.6. Acid leachable multielement determination by ICP-AES

The acid leachable metal content of soil samples was determined by the same method described above for sediment samples (refer to Section §6.3.3.2.5).

##### 6.3.3.3.7. Selection of residual soil data

A wide range of total multielement data determined by both ICP-MS and XRFS methods were available. Data quality, including results of CRMs and ANOVA data for duplicate analyses, was assessed on an element-by-element basis. If only one of the two techniques were sufficiently sensitive to permit the determination of a particular element at the background levels found in the samples, then this technique was automatically selected. If both techniques were sufficiently sensitive, the data from the technique with the best precision, based on the ANOVA duplicate analyses, were selected. The actual method used for each element is given in Table 6.4.

The quality of the ANOVA data for the ICP-AES analysis of the aqua regia solutions was also assessed. For a few elements, the analytical variation exceeded that of the geochemical variation, and the data for these elements were not considered suitable for publication. Those elements with acceptable data quality are given in Table 6.4. Aqua regia data were also used for total S, for which no data were available by any other technique.

#### 6.3.3.4. Analysis of humus samples

##### 6.3.3.4.1. Determination of mercury

Mercury in humus was determined by the method described above for sediments (refer to Section §6.3.3.2.2).

##### 6.3.3.4.2. Total multielement determination by ICP-MS

Humus samples were analysed for total element content by inductively coupled mass spectrometry. Two instruments were used: a Fisons Plasmaquad PQ2 ICP-MS during the early stages of the project and later an Argilent 7500i ICP-MS instrument. The calibration regime for both instruments was very similar.

Samples were extracted with nitric acid prior to analysis. Extraction with mild nitric acid was considered the most effective leaching method for environmental geochemical surveys to distinguish between elements, specifically Pb, Cd and Zn, derived from anthropogenic sources and elements derived from the natural background. The weak nitric acid digestion used was expected to dissolve mainly the adsorbed, exchangeable and carbonate-bound phases in the sample material.

Samples (400 mg of <0.063 mm) were placed in 50 ml polypropylene centrifuge tubes and 40 ml 4.5% sub-boiled HNO3 was added. The tubes were placed in an ultrasonic bath for 90 minutes and then centrifuged at room temperature at 3000 rpm for 30 minutes prior to analysis.

The estimated relative standard deviation was better than 5% for most of the anthropogenic sourced elements and always better than 10% for the elements reported. This estimate is valid, however, only for a well-homogenised reference sample; the sampling uncertainty was significant for the humus samples.

##### 6.4.3.4.3. Selection of humus data

The quality of the ANOVA data for Hg and the ICP-AES analysis of the aqua regia solutions was assessed. For a few elements, the analytical variation exceeded that of the geochemical variation, and these elements were not considered suitable for publication. Those elements with acceptable data quality are given in Table 6.4.

### 6.3.4. Quality assurance and control

Sampling and sample preparation are the most critical processes contributing to the overall uncertainty budget (Demetriades, 2014). The methods and quality control measures employed for sampling have been discussed earlier in this Manual (see Chapter 3). The methods employed for sample preparation are described in Chapter 4. A special homogeneity test protocol was established for the project to demonstrate the quality of sample preparation.

Approximately one per cent of the prepared samples were split into four sub-samples, two pressed pellets prepared from each sub-sample, and each pellet analysed twice by XRF for ten major and trace elements in order to control the ‘between-bottle’ homogeneity and ‘within-bottle’ homogeneity.

Uncertainty in chemical analysis, although often significantly less than that associated with sampling and sample preparation, is nevertheless essential to understand and control. All the participating laboratories listed in Table 6.2 were selected on the basis that they had established quality control systems in place, they regularly participated in international proficiency testing schemes for the methods for which they were responsible, and their estimated analytical uncertainty and sensitivity for all determinands was likely to be acceptable.

For the project, additional quality measures were used to provide extra assurance that the overall quality of data was up to the high standards required. For all solid sample analyses, two reference materials (ISE 921 and ISE 982) were analysed at regular intervals (between 1 and 2%, depending on the method) to monitor long-term stability and to enable the comparison of data from different methods and different laboratories. About 8% of the solid samples were also used for inter-laboratory comparison, e.g., about 50 soil samples were analysed by WD-XRF at BGS for comparison with the original BGR WD-XRF and GTK ICP-MS data.

To assess the precision of the chemical analysis relative to the variation due to sample preparation and sampling, duplicate samples were taken and repeat measurements were made.  Approximately 5% of all sites were sampled in duplicate. The duplicate samples were prepared independently of each other and analysed randomly along with all other samples, with each duplicate sample also analysed in duplicate to allow estimation of sampling uncertainty by the balanced ANOVA statistical interpretation method (Fig. 6.2).

Figure 6.2. FOREGS Atlas quality control of sampling and analyses by replicates using a balanced ANOVA design (modified from Sandström et al., 2005, Fig. 1, p.83).

### 6.3.5. Additional methods used in European continental-scale projects

To show the significance of collecting and analysing the Global Geochemical Reference Network samples in a short time to establish a homogeneous and comparable database worldwide, the analytical methods used in another two European multi-national continental-scale projects are described below, and their detection limits compared with those of the FOREGS Geochemical Atlas of Europe.

#### 6.3.5.1. Analytical methods used in the European Groundwater Geochemistry project

The material of this Section was mainly extracted from Birke et al. (2010), and Reimann and Birke (2010), and modified for the purposes of this Manual.

The reasons for selecting bottled mineral water to assess the geochemistry of European groundwater geochemistry are given in Reimann and Birke (2010).

The bottled mineral water samples were analysed by ICP-AES, ICP-QMS, IC, titration and photometric methods in the laboratories of the Bundesanstalt für Geowissenschaften und Rohstoffe (BGR) in Hannover Germany.

After adjusting the bottled mineral water samples to room temperature, pH (SenTix 81, glass electrode) and electrical conductivity (WTW TetraCon 325) were measured (WTW Multiline P4 electrode) with a precision of 0.01 for both parameters.

##### 6.3.5.1.1. ICP-QMS analyses

Sample preparation (acidification and transfer to polypropylene (PP) tubes) and the ICP-QMS analysis were carried out using cleanroom procedures (class 10,000 cleanroom) in the BGR laboratories. Disposable polyethylene (PE) gloves (all other gloves contain toxic elements, e.g., Zn) were used for all work in the cleanroom.

The water samples were analysed by inductively coupled plasma-quadrupole mass spectrometry (ICP-QMS) using an Agilent 7500ce instrument. The instrument is equipped with a standard peristaltic pump, a Micro Mist concentric nebuliser, a Peltier-cooled spray chamber, the Plasma Forward Power, the Shield Torch System, and a Collision/Reaction Cell system. An autosampler (ASX 520, Cetac Technologies) was employed to introduce the analytes into the plasma of the ICP-MS. Sample introduction was done with a Micro Mist concentric nebuliser and Scott-type quartz glass spray chamber.

Details of the instrument settings and scanning conditions are given in Table 6.5. The analyses were conducted according to the German norm DIN 38406-29 (DIN, 1996).

Hydrogen was considered to be useful as a reaction gas only for Se. Helium was used as the collision gas for Ag, As, Co, Cr, Cu, Fe, Ga, Ge, Mo, Ni, Sc, Ti and Zn. No cell gas was used for the other elements.

Normally the analysis of iodine by ICP-MS requires special sample treatment and elimination of polyatomic interferences, in addition to keeping iodine in solution. But, for bottled mineral water, which has weakly acidic to acid pH, iodine can be measured directly. For the determination of iodine in this study, however, the samples were acidified with 1% ultrapure HNO3. The instrument was calibrated with iodate. It was also attempted to stabilise iodine at basic pH, but the measurements were not reproducible.

Within the framework of the Geochemical Atlas of Germany (Birke et al., 2006), ten water samples from different sources were acidified with 1% ultrapure HNO3, and stored in fluorinated ethylene propylene (FEP) bottles for three years at low temperature (<4oC) and measured at different times. The measured iodine concentrations were reproducible over the entire three years. It has been reported that acidification can result, however, in the formation of elementary iodine, which can be adsorbed on the bottle walls and instrument tubing. But it has been verified that when HNO3 is used for acidification, the iodine is oxidised to the iodate; this can occur at different rates, and for this reason, there are no commercial certified standards for iodine. Water samples with high iodine concentrations were immediately reanalysed at several dilutions to test the stability of the measurement. After the measurement of a high iodine concentration, the water samples were reanalysed using the following procedure to exclude the possibility of a memory effect. A newly cleaned sample input system (tubing, sampler and skimmer cones, spray chamber) was used for each measurement. Each polypropylene (PP) tube was used only once.

Table 6.5. Instrument settings and scanning conditions of the Agilent 7500ce (ICP-MS) (from Birke et al., 2010, Table 2, p.220).

| Tuning Parameters |  |
| --- | --- |
| Plasma forward power | 1550 W |
| Reflected power | 1 W |
| Sampling depth | 8 mm |
| Cool gas flow | 15 l/min |
| Carrier gas flow | 0.72 l/min |
| Makeup gas flow | 0.28 l/min |
| He collision gas flow | 5.1 ml/min He 5.3 |
| H2 reaction gas flow | 4.2 ml/min H2 5.3 |
| Spray chamber | Scott/Duran |
| Spray chamber temp | 10°C |
| Nebuliser | MicroMist |
| Sample uptake | 0.4 ml/min |
| Oxide CeO156/Ce140 | 0.8 – 1.0 % |
| Doubly charged Ce140/70 | 1 – 3 % |
| Acquisition Parameters |  |
| Points per mass | 1 |
| Repetitions | 3 |
| Scans/repetition | 100 |
| Stabilisation sample | 60 sec |
| Stabilisation no gas | 15 sec |
| Stabilisation H2 | 15 sec |
| Stabilization He | 15 sec |
| Total acquisition time | 165 sec |
| Rinse time (intelligent) for Sb, Tl, U | 30 – 90 sec |

It is known that interference corrections must be used for the numerous mass overlaps of the various REE oxides. The collision cell conditions that were used cannot eliminate these overlaps.  A correction algorithm similar to that of Hall et al. (1995, 1996) was used to correct for the influence of BaO on the peaks for 151Eu and 153Eu. The correction was made for each analytical sample batch to take into consideration the rate of BaO formation (0.8–1.0 %). Appropriate settings for the Agilent mass spectrometer were selected so that REE oxide formation was minimised. It is known that many of the REEs require 3 to 4 such corrections to obtain accurate results.

Table 6.6 shows that the analytical methods used in the European Groundwater Geochemistry project (Birke et al., 2010; Reimann and Birke, 2010), approximately eight years after the FOREGS Geochemical Atlas of Europe project (Sandström et al., 2005), have lower detection limits, because of improvements in the technology of analytical instruments.

Table 6.6. Analytical methods and detection limits used in (a) the European Groundwater Geochemistry (EGG) project, and (b) Stream water of the FOREGS Geochemical Atlas of Europe. For the EGG project there are three detection limits, i.e., Instrument detection limit (IDL), Reported detection limit (RDL) and Practical detection limit (PDL); the latter was estimated by a modified Thompson & Howarth (1976, 1978) method (Demetriades, 2011; see Chapter 7 in this Manual). Colour notation:  Light Blue = lower detection limits (LDL) of European Groundwater Geochemistry project; Pale Brown = lower quantification limits of FOREGS Stream water; Pale Yellow = similar or better quantification limit of FOREGS Stream water (modified from Birke et al. (2010, Table 1, p.219-220) and Sandström et al. (2005, Table 3, p.89)).

| Parameter | Unit | (a) European Groundwater Geochemistry project |  |  |  |  |  |  | (b) FOREGS Atlas Stream water |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Analytical method | Isotope | Used gases | Instrument detection limit | Reported detection limit | Practical detection limit | Precision | Quantifi-cation limit | Analytical method |
|  |  |  |  |  | IDL | RDL | PDL | (%) | QL |  |
| Ag | µg/l | ICP-QMS | Min: 107-109 | He GMa | 0.001 | 0.002 | 0.002 | 13 | 0.002 | ICP-MS |
| Al | µg/l | ICP-QMS | 27 | NGMb | 0.2 | 0.5 | 0.2 | 5 | 0.1 | ICP-MS |
| As | µg/l | ICP-QMS | 75 | He GMa | 0.01 | 0.03 | 0.001 | 10 | 0.01 | ICP-MS |
| B | µg/l | ICP-QMS | 11 | NGMb | 0.1 | 2 | 0.2 | 4 | 0.01 | ICP-MS |
| Ba | µg/l | ICP-QMS | 137 | NGMb | 0.005 | 0.1 | 0.2e | 5 | 0.005 | ICP-MS |
| Be | µg/l | ICP-QMS | 9 | NGMb | 0.001 | 0.01 | 0.005 | 5 | 0.005 | ICP-MS |
| Bi | µg/l | ICP-QMS | 209 | NGMb | 0.0005 | 0.005 | 0.003 | /f | 0.002 | ICP-MS |
| Ca | mg/l | ICP-AES | / | / | 0.005 | 0.01 | / | / | 0.001 | ICP-AES |
| Cd | µg/l | ICP-QMS | 114_cor | NGMb | 0.001 | 0.003 | 0.002 | 29 | 0.002 | ICP-MS |
| Ce | µg/l | ICP-QMS | 140 | NGMb | 0.0005 | 0.001 | 0.0001 | 13 | 0.002 | ICP-MS |
| Co | µg/l | ICP-QMS | 59 | He GMa | 0.002 | 0.01 | 0.002 | 5 | 0.005 | ICP-MS |
| Cr | µg/l | ICP-QMS | 52 | He GMa | 0.01 | 0.2 | 0.014 | 7 | 0.01 | ICP-MS |
| Cs | µg/l | ICP-QMS | 133 | NGMb | 0.0005 | 0.002 | 0.004e | 3 | 0.002 | ICP-MS |
| Cu | µg/l | ICP-QMS | Min: 63-65 | He GMa | 0.01 | 0.1 | 0.029 | 2 | 0.005 | ICP-MS |
| Dy | µg/l | ICP-QMS | 163 | NGMb | 0.0002 | 0.001 | 0.001 | 16 | 0.002 | ICP-MS |
| Er | µg/l | ICP-QMS | 166 | NGMb | 0.0002 | 0.001 | 0.0004 | 13 | 0.002 | ICP-MS |
| Eu | µg/l | ICP-QMS | 151_kor | NGMb | 0.0002 | 0.001 | 0.001 | 18 | 0.002 | ICP-MS |
| Fe | µg/l | ICP-QMS | 56 | He GMa | 0.01 | 0.5 | 0.08 | 4 | 1 | ICP-MS |
| Ga | µg/l | ICP-QMS | 71 | He GMa | 0.0005 | 0.005 | 0.005 | 4 | 0.002 | ICP-MS |
| Gd | µg/l | ICP-QMS | 157 | NGMb | 0.0002 | 0.002 | 0.001 | 22 | 0.002 | ICP-MS |
| Ge | µg/l | ICP-QMS | 72 | He GMa | 0.005 | 0.03 | 0.02 | 6 | 0.005 | ICP-MS |
| Hf | µg/l | ICP-QMS | 178 | NGMb | 0.0001 | 0.002 | 0.001 | 28 | 0.002 | ICP-MS |
| Hg | ng/l | AFS | / | / | 1 | 5 | / | / | / |  |
| Ho | µg/l | ICP-QMS | 165 | NGMb | 0.0001 | 0.001 | 0.0003 | 19 | 0.002 | ICP-MS |
| I | µg/l | ICP-QMS | 127 | NGMb | 0.01 | 0.2 | 0.09 | 15 | 0.01 | ICP-MS |
| In | µg/l |  | / | / | / | / | / | / | 0.002 | ICP-MS |
| K | mg/l | ICP-AES | / | / | 0.05 | 0.1 | / | / | / | ICP-MS |
| La | µg/l | ICP-QMS | 139 | NGMb | 0.0001 | 0.001 | 0.001 | 9 | / | ICP-MS |
| Li | µg/l | ICP-QMS | 7 | NGMb | 0.01 | 0.2 | 0.14 | 5 | 0.005 | ICP-MS |
| Lu | µg/l | ICP-QMS | 175 | NGMb | 0.00005 | 0.001 | 0.0002 | 16 | 0.002 | ICP-MS |
| Mg | mg/l | ICP-AES | / | / | 0.005 | 0.01 | / | / | 0.001 | ICP-AES |
| Mn | µg/l | ICP-QMS | 55 | NGMb | 0.005 | 0.1 | 0.111 | 2 | 0.05 | ICP-MS |
| Mo | µg/l | ICP-QMS | 95 | He GMa | 0.001 | 0.02 | 0.017 | 4 | 0.002 | ICP-MS |
| Na | mg/l | ICP-AES | / | / | 0.02 | 0.1 | / | / | 0.005 | ICP-AES |
| Nb | µg/l | ICP-QMS | 93 | NGMb | 0.001 | 0.01 | 0.001 | 15 | 0.002 | ICP-MS |
| Nd | µg/l | ICP-QMS | 146 | NGMb | 0.0001 | 0.001 | 0.001 | 18 | 0.005 | ICP-MS |
| Ni | µg/l | ICP-QMS | 60 | He GMa | 0.005 | 0.02 | 0.01 | 4 | 0.005 | ICP-MS |
| P | mg/l | ICP-AES | / | / | 0.01 | 0.01 | / | / | / | / |
| Pb | µg/l | ICP-QMS | Sum: 206+207+208 | NGMb | 0.001 | 0.01 | 0.003 | 6 | 0.005 | ICP-MS |
| Pr | µg/l | ICP-QMS | 141 | NGMb | 0.00005 | 0.001 | 0.0002 | 15 | 0.002 | ICP-MS |
| Rb | µg/l | ICP-QMS | 85 | NGMb | 0.001 | 0.01 | 0.30e | 6 | 0.002 | ICP-MS |
| Sb | µg/l | ICP-QMS | 121 | NGMb | 0.001 | 0.01 | 0.0001 | 6 | 0.002 | ICP-MS |
| Sc | µg/l | ICP-QMS | 45 | He GMa | 0.001 | 0.02 | 0.001 | 23 | / |  |
| Se | µg/l | ICP-QMS | 78 | H2 GMc | 0.005 | 0.02 | 0.009 | 19 | 0.01 | ICP-MS |
| Si | mg/l | ICP-AES | / | / | 0.03 | 0.03 | / | / | 0.01 | ICP-AES |
| Sm | µg/l | ICP-QMS | 147 | NGMb | 0.0001 | 0.001 | 0.001 | 23 | 0.002 | ICP-MS |
| Sn | µg/l | ICP-QMS | 118 | NGMb | 0.001 | 0.02 | 0.003 | 12 | / | ICP-MS |
| Sr | mg/l | ICP-AES | / | NGMb | 0.001 | 0.001 | / | / | 0.001 | ICP-AES |
| Ta | µg/l | ICP-QMS | 181 | NGMb | 0.0005 | 0.005 | 0.002 | /f | 0.002 | ICP-MS |
| Tb | µg/l | ICP-QMS | 159 | NGMb | 0.00005 | 0.001 | 0.0002 | 23 | 0.002 | ICP-MS |
| Te | µg/l | ICP-QMS | 126 | NGMb | 0.001 | 0.03 | 0.01 | /f | 0.005 | ICP-MS |
| Th | µg/l | ICP-QMS | 232 | NGMb | 0.0001 | 0.001 | 0.0004 | 33 | 0.002 | ICP-MS |
| Ti | µg/l | ICP-QMS | Min: 47-49 | He GMa | 0.005 | 0.08 | 0.04 | 52 | 0.01 | ICP-MS |
| Tl | µg/l | ICP-QMS | 205 | NGMb | 0.0005 | 0.002 | 0.001 | 6 | 0.002 | ICP-MS |
| Tm | µg/l | ICP-QMS | 169 | NGMb | 0.00005 | 0.001 | 0.0002 | 22 | 0.002 | ICP-MS |
| U | µg/l | ICP-QMS | 238 | NGMb | 0.00005 | 0.001 | 0.006e | 2 | 0.002 | ICP-MS |
| V | µg/l | ICP-QMS | 51 | NGMb | 0.01 | 0.1 | 0.02 | 6 | 0.05 | ICP-MS |
| W | µg/l | ICP-QMS | 182 | NGMb | 0.002 | 0.05 | 0.009 | 1 | 0.002 | ICP-MS |
| Y | µg/l | ICP-QMS | 89 | NGMb | 0.00005 | 0.001 | 0.002 | 7 | 0.002 | ICP-MS |
| Yb | µg/l | ICP-QMS | 172 | NGMb | 0.0001 | 0.001 | 0.001 | 17 | 0.002 | ICP-MS |
| Zn | µg/l | ICP-QMS | Min: 66-68 | He GMa | 0.01 | 0.2 | 0.11 | 3 | 0.01 | ICP-MS |
| Zr | µg/l | ICP-QMS | 90 | NGMb | 0.0001 | 0.001 | 0.013 | 7 | 0.002 | ICP-MS |
| Brˉ | mg/l | IC | / | / | 0.003 | 0.003 | / | / | 0.01 | IC |
| HCO3ˉ | mg/l | titration | / | / | 2 | 2 | / | / | / | / |
| Alkalinity | mg/l | titration | / | / | 0.1 | 0.1 | / | / | / | / |
| Clˉ | mg/l | IC | / | / | 0.01 | 0.01 | / | / | 0.1 | IC |
| Fˉ | mg/l | IC | / | / | 0.003 | 0.003 | / | / | 0.05 | IC |
| NH4+ | mg/l | photometric | / | / | 0.005 | 0.005 | / | / | / | / |
| NO2ˉ | mg/l | IC | / | / | 0.005 | 0.005 – 0.1d | / | / | / | / |
| NO3ˉ | mg/l | IC | / | / | 0.01 | 0.01 – 1.0d | / | / | 0.04 | IC |
| PO43ˉ | mg/l | ICP-AES | / | / | 0.02 | 0.02 | / | / | / | / |
| SO42ˉ | mg/l | IC | / | / | 0.01 | 0.01 | / | / | 0.3 | IC |
| SiO2 | mg/l | ICP-AES | / | / | 0.05 | 0.05 | / | / | / | / |

Superscript notation: aHe gas mode; bno gas mode; cH2 gas mode; ddependent on TDS; etoo large a concentration range and too many high values to reliably calculate PDL; finsufficient number of values above detection limit to estimate precision.

##### 6.3.5.1.2. ICP-AES analyses

The acidified and non-acidified water samples were analysed using an ICP-AES (Spectro Ciros instrument) with radial plasma measurement according to DIN EN ISO 11885 (DIN, 1988, 1998, 2009). The wavelengths and detection limits are given in Table 6.7 for the analysed elements. The detection limits were estimated from the calibration curves according to DIN 32645 (DIN, 1994).

Blank samples (0.15 M HNO3 in deionised water) were measured for calibration of the instrument. A multielement standard solution containing the elements of Table 6.7 in 0.15 M HNO3 was also used. This standard of known composition was analysed as part of each batch of samples.

##### 6.3.5.1.3. Ion chromatography (IC)

A Dionex ICS 3000 ion chromatograph, equipped with a conductivity detector, an AS 19 column and a self-regenerating suppressor using a KOH eluent, was used for the determination of F-, Cl-, NO2-, Br-, NO3- and SO42- in nonacidified water samples, according to DIN EN ISO 10304-1 (DIN, 1995). The detection limits of this method are given in Table 6.6.

##### 6.3.5.1.4. Photometric analyses

Ammonium ion was analysed photometrically using a Thermo Unicam UV 300 spectrometer according to DIN 38406-5 (DIN, 1983) in samples that were not acidified. The detection limits estimated from the calibration curve, and calculated according to DIN 32645 (DIN, 1994), are given in Table 6.6.

Table 6.7. Wavelengths and instrument detection limits for elements analysed in groundwater samples by ICP-AES (from Birke et al., 2010, Table 3, p.220).

| Element | Wavelength (nm) | Instrument detection limit (mg/l) |
| --- | --- | --- |
| Al | 167.078 | 0.003 |
| B | 249.678 | 0.01 as BO2 |
| Ba | 455.404 | 0.0005 |
| Be | 313.042 | 0.0005 |
| Ca | 317.933, 422.673 | 0.005 |
| Cd | 228.802 | 0.002 |
| Co | 228.616 | 0.003 |
| Cr | 267.716 | 0.003 |
| Cu | 324.754 | 0.003 |
| Fe | 259.94 | 0.003 |
| K | 766.49 | 0.05 |
| Li | 670.78 | 0.003 |
| Mg | 285.213, 279.079 | 0.005 |
| Mn | 257.611 | 0.001 |
| Na | 589.592 | 0.02 |
| Ni | 231.603 | 0.005 |
| P | 177.49 | 0.01 as PO43- |
| Pb | 220.351 | 0.02 |
| Sc | 361.364 | 0.001 |
| Si | 251.611 | 0.03 as SiO2 |
| Sr | 407.771 | 0.0003 |
| V | 311.071 | 0.003 |
| Zn | 213.856 | 0.003 |

##### 6.3.5.1.5. Titration method

Alkalinity was determined on nonacidified water samples by titration with 0.05 M HCl. This was done with an automatic Metrohm titrator on a 50 ml water sample. The titration result was given as bicarbonate concentration (titration end-point 4.3 in acid pH range) or carbonate (titration end-point 8.2 in weakly basic pH range). The detection limit was 1 mg HCO3-/l.

##### 6.3.5.1.6. Atomic fluorescence spectrometry (AFS)

Atomic fluorescence spectrometry was used for the determination of mercury (Instrument PSA 10.035 Millennium Merlin 1631).

Elemental mercury vapour is generated from the sample by reduction with tin(II) chloride, and is purged from the solution by an argon carrier stream. The mercury vapour is detected by atomic fluorescence spectrometry.

The system uses continuous flow vapour generation, the mercury vapour is passed over a gold trap where the mercury is concentrated. After collection for a set time (25 sec), the mercury is then thermally desorbed from the gold into the atomic fluorescence detector.

The sample is stabilised by a mixture of potassium dichromate and nitric acid (5 g potassium dichromate in 1 litre HNO3 (64%), 1 ml in 100 ml sample).

The linear dynamic range of this method is approximately 0.1 ng/l to 100 µg/l Hg.

Accuracy and precision were controlled continuously by using CRMs BCR 60 and BCR 61.

High purity MERCK reagents (suprapur) were used in all cases; prior to use, the tin(II) solution was purged for 24 hours with nitrogen.

#### 6.3.5.2. Analytical methods used in the European Geochemical Mapping of Agricultural and grazing land Soil project

The material of this Section was extracted from reports of the GEMAS project (Geochemical Mapping of Agricultural and grazing land Soils), namely Reimann et al. (2009, 2011) and Birke et al. (2014), and modified for the purposes of this Manual.

##### 6.3.5.2.1. Total element concentrations by XRF

The elements SiO2, TiO2, Al2O3, Fe2O3, MnO, MgO, CaO, Na2O, K2O, P2O5, SO3, Cl, F, As, Ba, Bi, Ce, Co, Cr, Cs, Cu, Ga, Hf, La, Mo, Nb, Ni, Pb, Rb, Sb, Sc, Sn, Sr, Ta, Th, U, V, W, Y, Zn and Zr were determined by wavelength dispersive X-ray fluorescence spectrometry (WD-XRFS) using PANalytical PW2400 and AXIOS WD-XRFs, with Cr and Rh anode X-ray tubes, respectively.

The splits of all <2 mm samples were prepared for XRF analysis via milling to <0.063 mm particle size in a disk mill using agate vessels. Loss on ignition (LOI) was determined on all samples by slowly heating them to 1030°C and keeping the samples at this temperature for 15 minutes in a muffle furnace. Although LOI may be used as a proxy for TOC, it must be recognised that other components may be released, including free water, bound water, carbonate minerals, sulphate and sulphides. Depending on the value of LOI (above or below 25%), the samples for XRF determination were prepared, using different proportions of lithium metaborate and lithium bromide for the fused beads. The sample weight in each case was 1000 mg.

A 1000 mg aliquot of each soil sample with a LOI< 25% was mixed with 5.0 g lithium metaborate and 25 mg lithium bromide in Pt95-Au5 crucibles, and fused at 1200°C for 20 minutes in an automatic fluxer (HAG 12-1500). For soil samples with a LOI>25%, weights of 2.5 g lithium metaborate and 2.415 g lithium tetraborate were used for the preparation of the fused beads. To correct for matrix and spectral interferences, calibration curves were constructed using 130 certified reference materials.

XRF is the method of choice when total element concentrations in a sample are to be determined. However, very light elements (like Li, Be, B), or trace elements with very low concentrations (like Ag, Au, Se, Te) cannot be reliably determined by XRF (Reimann et al., 2011).

Table 6.8 shows that the XRF instrument of BGR, used in the analysis of the GEMAS agricultural soil samples, has comparable detection limits to those of the FOREGS Geochemical Atlas of Europe. However, the estimation of the practical detection limit in the GEMAS project indicates that lower detection limits can be used. As the XRF results are censored, the estimated practical detection limits, in this case, verify the laboratory detection limits, and lower values cannot be assigned. Hence, it is strongly recommended for the XRF instruments to be programmed to provide uncensored values.

##### 6.3.5.2.2. Hot aqua regia extraction and determination by ICP-AES & ICP-MS

A weight of 15 g of the sieved mineral soil samples (<2 mm) was digested in 90 ml aqua regia and leached for one hour in a hot (95°C) water bath. After cooling, the solution was made up to a final volume of 300 ml with 5% HCl acid. The sample weight to solution volume ratio is 1 g per 20 ml. The solutions were then analysed using a Spectro Ciros Vision emission spectrometer (ICP-AES) and a Perkin Elmer Elan 6000/9000 inductively coupled plasma mass spectrometer (ICP-MS). In addition to the project standards and duplicate sample splits that were unknown to the analyst, the laboratory inserted its own project standards and analytical replicates.

The hot aqua regia digestion method, used in the GEMAS project, owing to the combined use of modern ICP-MS and ICP-AES instruments, and the larger weight of 15 g soil compared to the 1 g used in the FOREGS project, achieved lower detection limits after approximately eight years from the FOREGS Geochemical Atlas of Europe project in which an ICP-AES instrument was used (Table 6.9). It is, therefore, strongly recommended for the aqua regia extraction to use a larger weight.

Table 6.8. XRF methods and detection limits used in (a) the European Agricultural and Grazing Soil (GEMAS) project, and (b) FOREGS Geochemical Atlas of Europe determined at Bundesanstalt für Geowissenschaften und Rohstoffe (BGR) and British Geological Survey (BGS). For the GEMAS project, two detection limits are given, i.e., Instrument detection limit (IDL) and Practical detection limit (PDL); the latter was estimated by a modified Thompson & Howarth (1976, 1978) method (Demetriades, 2011; see Chapter 7 in this Manual). For the FOREGS project the Quantification limit (QL), and Upper Continental Crust abundance values are given for comparison. The precision at the 95% confidence level for the GEMAS project samples was estimated by two different methods. Colour notation: Light Blue = lower detection limits of GEMAS project soil samples analysed by XRF at BGR; Pale Brown = lower quantification limits of FOREGS sediment samples analysed by XRF at the British Geological Survey (BGS); Pale Yellow = similar or better quantification limit of FOREGS residual soil samples analysed by XRF at BGR (modified from Reimann et al. (2011, Table 2, p.19-20) and Demetriades et al. (2014, Table 6.1, p.50-51) and Sandström et al. (2005, Table 3, p.89)).

| Parameter | BGR Lab | (a) GEMAS Agricultural (Ap) soil |  |  | (b) FOREGS Geochemical Atlas of Europe |  | Upper Continental Crust |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Project replicates (N=104 pairs) |  |  |  |  |  |  |
|  |  | Practical | Precision % |  | Soil samples | Sediment samples |  |  |
|  | IDL | PDL | 1 | 2 | BGR QL | BGS QL | Abundance | Source |
| Values in wt.% |  |  |  |  |  |  |  |  |
| SiO2 | 0.1 | 0.01 | 1 | 0.4 | 0.1 | 0.1 | 66.6 | Rudnick and Gao (2003) |
| TiO2 | 0.001 | 0.005 | 1 | 2 | 0.001 | 0.005 | 0.640 | Rudnick and Gao (2003) |
| Al2O3 | 0.05 | 0.04 | 1 | 1 | 0.05 | 0.1 | 15.4 | Rudnick and Gao (2003) |
| Fe2O3 | 0.01 | 0.01 | 1 | 1 | 0.01 | 0.01 | 5.60 | Rudnick and Gao (2003) |
| MnO | 0.001 | 0.0001 | 2.2 | 2 | 0.001 | 0.01 | 0.100 | Rudnick and Gao (2003) |
| MgO | 0.01 | 0.01 | 1 | 1 | 0.01 | 0.1 | 5.19 | Rudnick and Gao (2003) |
| CaO | 0.005 | 0.01 | 2 | 1 | 0.01 | 0.05 | 3.59 | Rudnick and Gao (2003) |
| Na2O | 0.01 | 0.01 | 3 | 3 | 0.01 | 0.2 | 3.27 | Rudnick and Gao (2003) |
| K2O | 0.005 | 0.001 | 2 | 2 | / | / | 2.80 | Rudnick and Gao (2003) |
| P2O5 | 0.001 | 0.001 | 2 | 3 | 0.001 | 0.01 | 0.150 | Rudnick and Gao (2003) |
| SO3 | 0.01 | 0.003 | 12 | 27 | / | / | 0.015 | Rudnick and Gao (2003) |
| Cl | 0.002 | 0.002 | 43 | 67* | / | / | 0.037 | Rudnick and Gao (2003) |
| F | 0.05 | 0.04 | 50 | 341* | / | / | 0.056 | Rudnick and Gao (2003) |
| LOIa | 0.01 | 0.04 | 7 | 1 | / | / | / | / |
| Sb | 0.01 | 0.005 | 32 | 48 | / | / | 0.0062 | Rudnick and Gao (2003) |
| Cb | 0.01 | 0.008 | 22 | 6 | / | / | 0.324 | Wedepohl (1995) |
| TOCc | 0.1 | 0.3 | 22 | 9 | / | / | / | / |
| Values in mg/kg |  |  |  |  |  |  |  |  |
| As | 3 | 0.4 | 15 | 21 | / | 1 | 5.7 | Hu and Gao (2008) |
| Ba | 5 | 2 | 3 | 1 | 5 | 3 | 628 | Rudnick and Gao (2003) |
| Bi | 3 | 2 | 31 | 576* | / | / | 0.23 | Hu and Gao (2008) |
| Ce | 20 | 7 | 15 | 33 | / | / | 63 | Rudnick and Gao (2003) |
| Co | 3 | 1 | 16 | 27 | / | / | 15 | Hu and Gao (2008) |
| Cr | 4 | 3 | 3 | 3 | 3 | 3 | 73 | Hu and Gao (2008) |
| Cs | / | / | / | / | / | 4 | 4.9 | Hu and Gao (2008) |
| Cu | 5 | 0.5 | 10 | 21 | / | 1 | 27 | Hu and Gao (2008) |
| Ga | 2 | 1 | 8 | 20 | / | 1 | 18.6 | Hu and Gao (2008) |
| Hf | 5 | 9 | 37 | 92 | / | / | 5.3 | Rudnick and Gao (2003) |
| La | 20 | 64 | 27 | 120 | / | / | 31 | Rudnick and Gao (2003) |
| Mo | 2 | ** | 34 | ** | / | / | 0.6 | Hu and Gao (2008) |
| Nb | 2 | 0.2 | 8 | 15 | / | 1 | 11.6 | Hu and Gao (2008) |
| Ni | 3 | 1 | 4 | 5 | / | 1 | 34 | Hu and Gao (2008) |
| Pb | 3 | 0.3 | 98 | 13 | / | 1 | 17 | Rudnick and Gao (2003) |
| Rb | 2 | 0.2 | 4 | 6 | 2 | 1 | 94 | Hu and Gao (2008) |
| Sb | 5 | ** | 18 | ** | / | / | 0.75 | Hu and Gao (2008) |
| Sc | 2 | 1 | 10 | 33 | / | / | 14 | Hu and Gao (2008) |
| Sn | 4 | 2 | 45 | 577 | 2 | 1 | 2.2 | Hu and Gao (2008) |
| Sr | 2 | 1 | 1 | 1 | 2 | 1 | 320 | Rudnick and Gao (2003) |
| Ta | 5 | ** | 13 | ** | / | / | 0.92 | Hu and Gao (2008) |
| Th | 5 | 3 | 21 | 54 | / | 1 | 10.5 | Rudnick and Gao (2003) |
| U | 3 | 7 | 31 | 126 | / | 1 | 2.6 | Hu and Gao (2008) |
| V | 5 | 2 | 3 | 5 | / | 2 | 106 | Hu and Gao (2008) |
| W | 5 | ** | 22 | ** | 5 | 0.05 | 1.4 | Hu and Gao (2008) |
| Y | 3 | 2 | 4 | 1 | 3 | 0.02 | 21 | Rudnick and Gao (2003) |
| Zn | 3 | 1 | 2 | 2 | 3 | 1 | 75 | Hu and Gao (2008) |
| Zr | 3 | 1 | 3 | 4 | 3 | 1 | 193 | Rudnick and Gao (2003) |
| CECd | 0.1 | 0.04 | 6 | 9 | / | / | / | / |
| pH_CaCl2e | 0.1 | 0.1 | 1 | 3 | / | / | / | / |
| aGravimetric method; bLECO; cInfrared method; dAg thiourea method in meq/100g; epH-meter (Reimann et al., 2011; Birke et al., 2014) |  |  |  |  |  |  |  |  |
| * = too many values near DL to estimate reliable values of PDL and Precision |  |  |  |  |  |  |  |  |
| ** = too few samples above detection to estimate PDL and Precision 2 |  |  |  |  |  |  |  |  |
| DL = Detection limit |  |  |  |  |  |  |  |  |
| QL = Quantification limit |  |  |  |  |  |  |  |  |

Table 6.9. Hot aqua regia digestion methods and determination of element concentrations by ICP-MS and ICP-AES and detection limits used in (a) the Geochemical Mapping of Agricultural and Grazing Soil (GEMAS) project, and (b) the FOREGS Geochemical Atlas of Europe. For the GEMAS project, two detection limits are given, i.e., Instrument detection limit (IDL) and Practical detection limit (PDL); the latter was estimated by a modified Thompson & Howarth (1976, 1978) method (Demetriades, 2011; see Chapter 7 in this Manual), and for the FOREGS project the Quantification limit (QL). Upper Continental Crust abundance values are given for comparison. The precision at the 95% confidence level for the GEMAS project samples has been estimated by two different methods. Colour notation: Light Blue = lower detection limits of GEMAS project soil samples analysed by a commercial laboratory; Pale Brown = lower quantification limits of FOREGS residual soil samples (modified from Reimann et al. (2009, Table 3, p.22-24) and Demetriades et al. (2014, Table 6.1, p.50-51) and Sandström et al. (2005, Table 3, p.89)).

| Element | (a) GEMAS Agricultural soil (Ap) samples |  |  |  |  |  | (b) FOREGS Residual soil samples | Upper Continental Crust |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Commercial Laboratory | Lab replicates (N=146 pairs) | Project replicates (N=104 pairs) |  |  | Lab replicates (N=146 pairs) |  |  |  |
|  |  | Practical | Practical | Precision % |  | Precision % |  |  |  |
|  | IDL | PDL | PDL | 1 | 2 | 1 | Lab QL | Abundance | Source |
| Ag | 0.002 | 0.002 | 0.007 | 9 | 18 | 8 | / | 0.053 | Rudnick and Gao (2003) |
| Al | 100 | 104 | 4 | 6 | 8 | 3 | / | 81,505 | Rudnick and Gao (2003) |
| As | 0.1 | 0.065 | 0.005 | 8 | 10 | 4 | 5 | 5.7 | Hu and Gao (2008) |
| Au | 0.0002 | 0.002 | 0.0007 | 46 | 55 | 66 | / | 0.0015 | Rudnick and Gao (2003) |
| Ba | 0.5 | 0.88 | 1.3 | 7 | 8 | 5 | 1 | 628 | Rudnick and Gao (2003) |
| Be | 0.1 | 0.04 | 0.033 | 19 | 39 | 16 | / | 1.9 | Hu and Gao (2008) |
| Bi | 0.02 | 0.004 | 0.01 | 6 | 7 | 5 | / | 0.23 | Hu and Gao (2008) |
| Ca | 100 | 48 | 178 | 11 | 4 | 6 | / | 25,657 | Rudnick and Gao (2003) |
| Cd | 0.01 | 0.01 | 0.0086 | 8 | 9 | 9 | / | 0.06 | Hu and Gao (2008) |
| Ce | 0.1 | 1.21 | 1.2 | 6 | 14 | 4 | / | 63 | Rudnick and Gao (2003) |
| Co | 0.1 | 0.013 | 0.085 | 6 | 8 | 4 | 1 | 15 | Hu and Gao (2008) |
| Cr | 0.5 | 0.45 | 0.019 | 6 | 8 | 5 | 1 | 73 | Hu and Gao (2008) |
| Cs | 0.02 | 0.014 | 0.011 | 7 | 9 | 4 | / | 4.9 | Hu and Gao (2008) |
| Cu | 0.01 | 0.02 | 0.27 | 5 | 4 | 4 | 1 | 27 | Hu and Gao (2008) |
| Fe | 100 | 121 | 163 | 3 | 5 | 3 | 2 | 39,176 | Rudnick and Gao (2003) |
| Ga | 0.1 | 0.026 | 0.065 | 5 | 6 | 4 | / | 18.6 | Hu and Gao (2008) |
| Ge | 0.1 | 0.64 | 0.019 | 52 | 58 | 47 | / | 1.3 | Hu and Gao (2008) |
| Hf | 0.02 | 0.006 | 0.01 | 12 | 9.5 | 14 | / | 5.3 | Rudnick and Gao (2003) |
| Hg | 0.005 | 0.003 | 0.002 | 13 | 15 | 13 | / | 0.05 | Rudnick and Gao (2003) |
| In | 0.02 | 0.003 | 0.004 | 21 | 28 | 19 | / | 0.066 | Hu and Gao (2008) |
| K | 100 | 5.7 | 21 | 8 | 13 | 5 | / | 23,244 | Rudnick and Gao (2003) |
| La | 0.5 | 0.09 | 0.026 | 5 | 8.5 | 4 | / | 31 | Rudnick and Gao (2003) |
| Li | 0.1 | 0.04 | 0.28 | 6 | 8 | 4 | / | 41 | Hu and Gao (2008) |
| Mg | 100 | 3.1 | 64 | 5 | 6 | 3 | / | 14,955 | Rudnick and Gao (2003) |
| Mn | 1 | 14 | 2 | 6 | 7 | 8 | 10 | 774 | Rudnick and Gao (2003) |
| Mo | 0.01 | 0.01 | 0.018 | 9 | 17 | 7 | / | 0.6 | Hu and Gao (2008) |
| Na | 10 | 1.1 | 1.1 | 8 | 14 | 10 | / | 24,259 | Rudnick and Gao (2003) |
| Nb | 0.02 | 0.01 | 0.028 | 11 | 9 | 7 | / | 11.6 | Hu and Gao (2008) |
| Ni | 0.1 | 0.11 | 0.26 | 6 | 6.5 | 6 | 2 | 34 | Hu and Gao (2008) |
| P | 10 | 33 | 4.3 | 7 | 11 | 5 | / | 655 | Rudnick and Gao (2003) |
| Pb | 0.01 | 0.35 | 0.88 | 11 | 15 | 12 | 3 | 17 | Rudnick and Gao (2003) |
| Pd | 0.01 | 0.01 | 0.001 | 540 | 228 | 415 | / | 0.000526 | Park et al. (2012) |
| Pt | 0.002 | 0.0016 | 0.00007 | 131 | 89 | 76 | / | 0.000599 | Park et al. (2012) |
| Rb | 0.1 | 0.15 | 0.23 | 6 | 10 | 4 | / | 94 | Hu and Gao (2008) |
| Re | 0.001 | 0.0003 | 0.06 | 1079 | 100 | 1048 | / | 0.000198 | Rudnick and Gao (2003) |
| S | 200 | 4 | 3.3 | 8 | 7 | 46 | 50 | 62 | Rudnick and Gao (2003) |
| Sb | 0.02 | 0.003 | 0.0002 | 12 | 18 | 7 | / | 0.75 | Hu and Gao (2008) |
| Sc | 0.1 | 0.0006 | 0.05 | 5 | 7.5 | 5 | / | 14 | Hu and Gao (2008) |
| Se | 0.1 | 0.04 | 0.1 | 19 | 17 | 18 | / | 0.09 | Rudnick and Gao (2003) |
| Sn | 0.1 | 0.01 | 0.0056 | 8 | 14 | 20 | / | 2.2 | Hu and Gao (2008) |
| Sr | 0.5 | 0.4 | 0.65 | 9 | 6 | 8 | / | 320 | Rudnick and Gao (2003) |
| Ta | 0.05 | 0.005 | 0.004 | 65 | 84 | 73 | / | 0.92 | Hu and Gao (2008) |
| Te | 0.02 | 0.01 | 0.021 | 53 | 24 | 60 | / | 0.027 | Hu and Gao (2008) |
| Th | 0.1 | 0.02 | 0.0008 | 5 | 9 | 5 | / | 10.5 | Rudnick and Gao (2003) |
| Ti | 10 | 0.9 | 6.7 | 8 | 8 | 11 | / | 3836 | Rudnick and Gao (2003) |
| Tl | 0.02 | 0.0005 | 0.00009 | 9 | 11 | 10 | / | 0.55 | Hu and Gao (2008) |
| U | 0.1 | 0.04 | 0.015 | 5 | 5 | 5 | / | 2.6 | Hu and Gao (2008) |
| V | 2 | 0.09 | 0.07 | 8 | 10 | 5 | 1 | 106 | Hu and Gao (2008) |
| W | 0.1 | 0.01 | 0.0045 | 16 | 26 | 12 | / | 1.4 | Hu and Gao (2008) |
| Y | 0.01 | 0.04 | 0.053 | 4 | 9 | 3 | / | 21 | Rudnick and Gao (2003) |
| Zn | 0.1 | 0.26 | 1 | 4 | 5 | 5 | 1 | 75 | Hu and Gao (2008) |
| Zr | 0.1 | 0.02 | 0.07 | 12 | 8 | 8 | / | 193 | Rudnick and Gao (2003) |

## 6.4. Analysis of rock samples

Most routine chemical bulk analyses of rocks are currently made with instruments (the X-ray fluorescence spectrometer, and the inductively coupled plasma spectrometer), which permit the simultaneous or rapid successive determination of major and trace elements (Hooper, 1964, 1966; Welday et al., 1964; Hooper et al., 1976; Hall, 1992; Jarvis et al., 1992; Montaser, 1998; Olesik, 2000; Gäbler, 2002).

X-ray fluorescence spectrometry (XRF) coupled with total acid digestion and inductively coupled plasma-mass spectrometry (ICP-MS) analysis is the technique of choice of most petrologists and geochemists for whole-rock analysis. The XRF method is described in detail in the FOREGS (Sandström et al., 2005) and GEMAS atlases (Reimann et al., 2011; Birke et al., 2014) – refer to Sections §6.3.3.2.3, §6.3.3.3.4 and §6.3.5.2.1, respectively.

Fusion of a powdered sample with a suitable flux, usually lithium metaborate, lithium tetraborate or a blend of the two, is used to break down the mineral structures thereby eliminating sample heterogeneity, such as grain size and mineralogical effects of the various rock-forming minerals. Traditionally fusion methods have been predominantly used to determine the major elements in silicate rock, and pressed powder pellet is used for trace element analysis. More recently, the use of a single low dilution (2:1 flux:sample) fused bead has improved sensitivity allowing the analysis of major and trace elements at the same time (Johnson et al., 1999).

As detection limits of the XRF technique are often higher for the reliable analysis of many trace elements, it is desirable, as already pointed out, to couple the analysis of the same powdered sample by XRF and ICP-MS. Unless Li and B are of interest, the same fusion technique can be used to break down mineral structures making for a much more precise determination of trace elements, which are often located in refractory and hard-to-dissolve minerals such as zircon. The resulting fused bead is then re-ground and dissolved in strong acid in the usual way before analysis on the ICP-MS instrument. If Li and B are of interest, the powdered sample can be dissolved directly in a blend of strong inorganic acids before analysis. This approach, however, often fails to dissolve the most refractory minerals, leading to the name ‘near-total digestion’ technique and, therefore, misses some of the important trace elements such as Zr and the REEs. If boron is not of interest, a sodium borate flux can be used to fuse the sample (Bank et al., 2016).

Recent developments in whole-rock analysis have focused on using laser ablation to introduce the sample into the ICP-MS directly from the same fused bead as analysed by XRF (Eggins, 2003; Yu et al., 2003; Kurosawa et al., 2007). This approach reliably achieves complete sample digestion and provides for complementary XRF and LA-ICP-MS analysis of a full suite of major and trace elements including the rare earth elements (REEs) from single sample preparation (Fedorowich et al., 1993; Strnad et al., 2007; Tanaka et al., 2007; Liu et al., 2013). This reduces sample supply, preparation time and handling, thereby decreasing the risk of error in the analytical routine (e.g., Yu et al., 2003; Yamasaki and Yamashita, 2016). While this approach is still developing, this will become a significant method in the future.

Some specially adapted techniques need to be applied to elements that are difficult to analyse with either XRF or bulk ICP-MS, for example, noble elements (Au), platinum group elements (PGM), mercury, carbon, sulphur and halogens.

Modern instrumental methods such as flame atomic absorption spectrometry (F-AAS) (Terashima, 1988; Tewari et al., 1990; Ramesh et al., 2001; Balaram et al., 2012), graphite furnace atomic absorption spectrometry (GF-AAS) (Benedetti et al., 1987), neutron activation analysis (INAA) (Keays and Scott, 1976; Nesbitt et al., 1987), inductively coupled plasma atomic emission spectrometry (ICP-AES) (Yan et al., 1995), and inductively coupled plasma mass spectrometry (ICP-MS) (Hall and Bonham-Carter, 1988; Ely et al., 1999; Pitcairn et al., 2006) are commonly used for gold determination with low detection limits. The very low average gold concentration in non-mineralised rocks ranges between 0.0005 and 0.005 mg/kg and requires a separation (dissolution of gold from the geological samples), pre-concentration of gold and a correspondingly highly sensitive analytical method for its quantification (Pitcairn et al., 2006). Accordingly, the determination of gold in rock samples with concentrations in the ultra-trace range (in the parts-per-trillion to parts-per-billion) is considerably more complex than in ore samples where fire-assay (FA) and cyanidation methods have been successfully used for centuries (Chow and Beamish, 1967; Vanloon and Barefoot, 1991; Marsden and House, 2006; Asamoah et al., 2014).

Platinum-group elements (PGEs: Pt, Pd, Rh, Ir, Os, Ru) in rocks are best determined by fire assay with ICP-MS or HR-ICP-MS finish. Modern methods such as INAA and ICP-MS (including HR-ICP-MS) offer highly sensitive determinations for PGEs (Shazali et al., 1987; Hall and Pelchat, 1994; Figueiredo et al., 2000; Garuti et al., 2000; Hann et al., 2001; Morcelli et al., 2004; Balaram et al., 2006). Several geochemical studies have confirmed that fire assay with either ICP-MS or HR-ICP-MS is the best method for the determination of PGEs in geological samples (Jackson et al., 1990; Kontas et al., 1990; Flight et al., 1994; Barfoot, 1998; Reddi and Rao, 1999).

The precise and accurate determination of REEs in geological materials is mainly achieved by using ICP-MS-based methods with multi-collector inductively coupled plasma mass spectrometer (MC-ICP-MS) and isotope dilution thermal ionisation mass spectrometry (ID-TIMS). These techniques are superior in terms of analytical reproducibility compared with quadrupole ICP-MS (Baker et al., 2002; Kent et al., 2007). Other instrumental methods that are commonly used for the determination of REE are X-ray fluorescence spectrometry (XRF), instrumental neutron activation analysis (INAA) (Ravisankar et al., 2006; Bounouira et al., 2007; Alharbi and El-Taher, 2016) and HR-ICP-MS (Bradshaw et al., 1989; Smirnova et al., 2010; Balaram, 2018; Satyanarayanan et al., 2018; Singh et al., 2018).

Mercury (Hg) determination in rocks requires low-temperature aqua regia digestion followed by ICP-AES or ICP-MS finish with as low as a 1 mg/kg detection limit. For more than 40 years, total mercury has also been determined with a high level of sensitivity (detection limit: 0.005 mg/kg) directly on the rock powder (fine milled <0.04 mm) samples using flameless AAS (see Section §6.3.3.2.2).

An advantage of the XRF method is the measurement of halogens like chlorine, which is not possible with the classical ICP-MS technique. Otherwise, the analysis of halogens requires fusion techniques that allow the elements to be retained in solution, for example, KOH fusion followed by ion chromatography for fluorine analysis. In general, rocks can be analysed for fluorine and chlorine with a number of methods like ion-specific electrode (Rice, 1988; Yuchi et al., 1988), ion-chromatography (Hall et al., 1986; Kennedy et al., 1983; Gent and Wilson, 1985), XRF (Langenauer et al., 1992) or spectrophotometry (Fuge and Andrews, 1985; Hoffstetter et al., 1991). Bromine and iodine can be determined by spectrophotometry (Fuge et al., 1978); XRF (Ullmann and Tisue, 1983; Takeda et al., 2011) or instrumental neutron activation analysis (Langenauer et al., 1992).

Total carbon and sulphur can be analysed using infrared spectroscopy (Terashima, 1988). Automated, high-temperature combustion, carbon-sulphur-analysers utilise solid-state infrared absorption for the detection of total carbon and sulphur. In addition, secondary ion mass spectrometry (SIMS), nuclear methods (INAA, RNAA), and LA-ICP-MS can be used for the detection of low levels of sulphur in minerals and glasses (Ripley et al., 2011).

Summing up, the following bulk analytical methods are recommended to obtain the basic set of major and trace elements in rocks for the purpose of geochemical mapping: XRF, lithium (meta)borate (LiBO2) flux-fusion by ICP-MS, and FA by ICP-AES or ICP-MS for PGEs.

## 6.5. Determination of mineralogy by XRD

X-ray diffraction (XRD) is a powerful non-destructive technique for characterising all kinds of matter – ranging from fluids to powders and crystals. X-ray diffraction analysis is used, not only to identify the phase of an unknown substance and to estimate the lattice parameters, but also to determine the concentration of that phase in the mixture. The peak profile is also employed to estimate the particle size of very small crystals (crystallites) in a powder sample. Powder diffraction is one of the principal research tools of mineralogists since many minerals are available in polycrystalline form (Dinnebier and Billinge, 2008; Pecharsky and Zavalij, 2009; Lavina et al., 2014).

Qualitative XRD analysis usually involves the identification of phases in a specimen by comparison with ‘standard’ patterns (e.g., ICDD database, International Centre for Diffraction Data), and relative estimation of proportions of different phases in multiphase specimens by comparing peak intensities attributed to the identified phases (Ufer et al., 2008a, 2008b, 2012a, 2012b; Waseda et al., 2011; Lavina et al., 2014).

The most effective quantitative methods, particularly those involving pattern modelling, are computationally intensive, and can be applied only with powerful analytical software (e.g., GSAS, FullProf, FULLPAT, RockJock, Profex (Rietveld-Software BGMN), AUTOQUAN from GE Sensing & Inspection Technologies GmbH). The methods with the greatest chance of producing successful results generally involve the addition of a known amount of an internal standard and calculation of ratios of the areas of the standard peaks to those of the phases being determined.

The type and percentage of rock-forming minerals are characteristic of a rock type and are reflected in the content of the mineral-forming major, minor, trace and ultra-trace elements (Table 6.10). As soils and sediments are derived from the weathering and erosion of rocks, the determination of their mineralogy is important in the effective interpretation of geochemical patterns as has been shown by the “Geochemical and Mineralogical Maps for Soils of the Conterminous United States” (Smith et al., 2014). It is, therefore, strongly recommended that the Laboratory Committee includes in the analytical procedures the determination of mineralogy in all sample types.

Table 6.10. Estimated average concentrations of elements in the most common rock-forming minerals (slightly modified after Wedepohl, 1969-78). Major elements >1%; Minor elements >0.1 to <1%; Trace elements >0.01 to <0.1%; Ultra-trace elements <0.001%.

| Concentration | X% | 0.X% | 0.0X% | 0.00X% | 0.000X% |
| --- | --- | --- | --- | --- | --- |
| Minerals | Major elements | Minor elements | Trace elements | Ultra-trace elements |  |
| Plagioclase | Ca, Na, Si, Al | Sr, K | Ba, Rb, Ti, Mn | V, Zn, Ni | Pb, Cu, Li, Cr, Co |
| Alkali feldspar | K, Na, Si, Al | Ca, Ba, Sr | Rb, Ti | Pb, Li, Mn | Zn, V, Cr, Ni, Co |
| Quartz | Si |  |  | Al, Ti, Fe, Mg, Ca | Na, Li, Ni, Zn, Ge, Mn |
| Amphibole | Si, Fe, Na, Mg, Al, Ca | Ti, K, Mn, Rb | Zn, Cr, V, Sr, Ni | Ba, Cu, Co, Pb | Li, B |
| Pyroxene | Al, Si, Fe, Mg, Ca | Ti, Na, Mn, K | Cr, V, Ni, Sr | Cu, Co, Zn, Li, Rb | Ba, Pb, Ga, B |
| Biotite | Si, Al, Fe, Mg, K | Ca, Na, Ba, Mn, Rb | V, Cr, Li, Ni, Ti | Cu, Sr, Co, Pb | B |
| Olivine | Si, Fe, Mg | Ni, Cr, Ti, Ca | Mn, Co | Zn, V, Cu, Sc | Rb, B, Ge, Sr, As, Ga, Pb |
| Spinels (including magnetite) | Ti, Al, Fe | Mg, Mn, V | Cr, Zn, Cu | Ni, Co | Pb, Mo |

## 6.6. Particle size analysis

There are different methods for the determination of particle size (grain size) distribution (PSD) in soil and sediment samples. The Laboratory Committee should select the most appropriate, reproducible, cost- and time-effective method. The available and most used methods are:

- Köhn pipette method is described in ISO 11277 (ISO, 2020b).

- Hydrometer analysis method is described in ASTM D7928-16 (ASTM, 2017).

- Hydrophotometer analysis (Jordan et al., 1971).

- SediGraph particle size analysis (Coakley and Syvitski, 2010).

- Laser diffraction analysis (LDA) (Agrawal et al., 2010; ISO, 2020b), and

- Mid-infrared spectroscopy (MIRS) (Thomas et al., 2021).

The Köhn pipette method is a very time-consuming procedure and in practice impossible for very large sample sets. Alternatives to the Köhn-pipette method are the Hydrophotometer and the SediGraph particle size analysis. Comparability of SediGraph results to those of ISO 11277 (ISO, 2020b) has been demonstrated by Müller et al. (2009). Further, Coates and Hulse (1985) compared the first four methods of particle size analysis on fine-grained sediments and concluded that the results of the Köhn pipette and Hydrometer analysis methods are comparable and have excellent reproducibility. They also found that the results of the SediGraph compare well with those of the Köhn pipette and Hydrometer at particle sizes finer than 16 μm, but tend to read low at coarser sizes. However, the first four methods are time-consuming and not suitable when thousands of samples need to be analysed. Usually, laser-based methods are used in the case of thousands of samples. Nevertheless, it should be thoroughly tested with soil and sediment samples from different morphoclimatic regions, because the laser diffraction method is repeatable and characterised by a high-resolution. However, the measurement of soil particle size distribution can have limited reproducibility (Polakowski et al., 2021).

Another method that can be used in the case of thousands of samples is mid-infrared spectroscopy, the results of which can be used to predict, apart from the particle size distribution (Janik et al., 2011; Reimann et al., 2012; Birke et al., 2014), a large variety of soil properties, provided proper models for the soil and sediment samples under study are built and validated (Viscarra Rossell et al., 2006; Janik et al., 2009; Soriano-Disla et al., 2014; Janik et al., 2015a, b). The advantages and disadvantages of LDA and MIRS methods for particle size analysis of soil samples are discussed by Thomas et al. (2021).

Refer to Section §6.3.3.3.1 where the method used in the FOREGS Geochemical Atlas of Europe project for the determination of particle size distribution in soil samples is described.

## 6.7. Summary of recommendations

The analytical requirements for the purpose of producing internationally comparable geochemical baseline data are as follows:

- Given the many applications of geochemical baseline data, a comprehensive multi-element approach is essential. Analytical requirements are considered in two categories:

- For the Global Geochemical Reference Network samples, and

- National or regional surveys conducted by national or international organisations.

- Samples collected as part of the Global Geochemical Reference Network project should ultimately provide abundance data for at least 78 naturally occurring elements (excluding inert gases, except Rn, and some of the elements from the actinide series with atomic number >84, using methods with detection limits significantly below presently estimated crustal abundances.

- Analytical requirements for national or regional surveys must be fully compatible with those that will be used for the establishment of the Global Geochemical Reference Network database. Where possible identical methods should be used but alternative multi-element schemes, based on the techniques of XRF, ICP-AES, ICP-MS, INAA, etc., can be employed as long as the criteria of precision and accuracy are met, and the SRMs developed for the Global Geochemical Reference Network are included in the analytical batches (refer to Chapter 8 in this Manual). To portray the spatial distribution of elements as completely as possible, detection limits must be as low as possible.

- A proposal is made for a standard list of elements to be determined; elements are grouped into List 1 (51 elements) and List 2 (20 elements) (see Fig. 6.1); Ru, Rh, Re, Os and Ir require the establishment of satisfactory methods and detection limits; Ra and Rn require radiometric methods.

- For national geochemical surveys, if analytical facilities are insufficient to cover all List 1 elements initially, missing data should be added at a later date. List 2 is of lower priority, but potentially important.

- To establish reproducible geochemical baseline data of permanent value, analytical methods should be employed that provide the total concentration of each element present. Sample decomposition, where required, must, therefore, be total.

- Partial decomposition (partial extraction) methods have many variants in different laboratories and are difficult to standardise, so they are NOT recommended for any data that may be used for international compilation or correlation. The only exception is hot aqua regia extraction, which is the method used for the estimation of statutory values, and a standardised method should be used by all geological survey laboratories.

- The different options for the analysis of the Global Geochemical Reference Network samples are:

(a) All sample types are to be analysed in the same laboratory by the same analytical methods for the same suite of determinands. If this is not possible, then the second option is:

(b) In selected laboratories the same suite of determinands to be analysed by the same analytical methods on all samples of one sample type, i.e., rock, residual soil, humus, stream sediment, stream water, overbank sediment, and floodplain sediment.

In both cases, strict internal and external quality control procedures must be followed, and the inclusion in the analytical batches of SRMs, developed for the Global Geochemical Reference Network project. The selected laboratories must be approved by the relevant project management committees, namely the Laboratory and Quality Control Committees (see Chapter 10 in this Manual). The laboratories should be selected on the basis that they are accredited and have established quality control systems in place, they regularly participate in international proficiency testing schemes for the methods for which they are responsible, and their estimated analytical uncertainty and sensitivity for all determinands are likely to be acceptable.

- With respect to national and regional surveys, it is recommended that if more than 20% of the reportable values for any element determined fall below the detection limit, the results for that element should be considered unsatisfactory and alternative analytical methods should be considered.

- Strict quality control, through the use of appropriate primary (PRMs) and secondary (SRMs) reference materials, is essential, and how standard reference materials have been used and the resulting quality control statistics must be reported by each laboratory. The laboratory must provide a full report with analytical results of each sample suite in the order analysed, together with the analytical batch number, date of analysis, results of blanks, and repeated analyses of samples. It is important for the approved laboratory to have installed its own internal quality control procedures, and should include primary reference materials (PRMs) usually at a rate of four PRMs at the start and end of each analytical batch (see Chapter 8 in this Manual). The applied geochemist in charge of the Global Geochemical Reference Network project, as well as the national project managers, should set up their own independent or external quality control procedures in collaboration with the Quality Control Committee (see Chapter 10 in this Manual).

- The Canadian STSD and Chinese series GSD and GSS standard reference samples, are good reference materials for use as Primary Reference Materials (PRMs) in (a) international geochemical mapping, and (b) national geochemical surveys. However, for the Global Geochemical Reference Network project, it is considered imperative that large quantities of secondary reference materials (SRMs) for each sampling medium (rock, residual soil, stream sediment, overbank/floodplain sediment) should be developed. This is a mandatory prerequisite before the start of the project, because it is necessary to have adequate material not only until the completion of the global project, but there should be enough material for

(a) to be distributed to all countries for their regional geochemical surveys, and

(b) at least one repeat of the global project at a future date.

At least five SRMs of each sample medium with different element concentrations must be developed (refer to Chapter 8 in this Manual). For national surveys, primary (PRMs) and secondary reference materials (SRMs) should be used to monitor the analytical accuracy: PRMs to monitor the international or interlaboratory bias, and SRMs to monitor the routine between-batch bias within a laboratory.

- The laboratories must provide the uncensored values of determined elements, as these are important for the estimation of the practical detection limit of each determinand (refer to Chapter 7 in this Manual), which may be different from the reported laboratory detection limit.

- The laboratories must provide all the analytical results of each sample type, including the analytical reagent blanks, international and internal reference samples in the order analysed, together with the analytical batch numbers, date and time stamp for every analysis.

- The Project Manager in collaboration with the Laboratory Committee should organise the contract with the laboratory or laboratories, which will be participating in the analysis of all sample types collected in the Global Geochemical Reference Network project. The contract should specify that:

(a) Primary reference materials (PRMs) should be included at a rate of four PRMs at the start and end of each analytical batch.

(b) Replicated analyses to be performed every twentieth sample.

(c) Depending on the number of samples in each analytical batch 2 to 10 reagent blank samples should be included, and

(d) The uncensored values of determined elements and parameters MUST be provided, and the other information included in bullet 13 above, as well as any problems encountered, and solutions given.

An important condition in the contract is if the Quality Control Committee discovers any analytical errors in the results, the laboratory will be obliged to reanalyse the batches in question, and if all results are considered unacceptable the laboratory will be required to reanalyse the complete sample suite.

- Natural radioactivity of K, U, Th and Total is measured in the field using a portable threshold scintillometer, which should be calibrated at least once a year at a national facility, which is certified by the International Atomic Energy Agency.

Finally, it is strongly recommended that the Laboratory Committee consults the extensive list of analytical methods, which were published in the special issue of the Journal of Geochemical Exploration entitled “Geoanalysis” (Hall, 1992). This is still a reference book for applied geochemists on the state-of-the-art at that time. It describes not only the capabilities and limitations of major methodologies, but the accuracy and precision expected, practicalities of usage, the suite of international reference samples available and discussions on data quality.

Other publications on analytical methods that should be consulted are by Fletcher (1981), Hall (1991), Finch et al. (1992), Giessmann and Grebb (1994), Balaram (1995, 2018, 2020), Jenkins and Snyder (1996), USEPA (1996), Hall and Pelchat (1997), Odegard et al. (1998), CEN (2002), Gäbler (2002), Krachler et al. (2002, 2007), Yu et al. (2003), Townsend and Poole (2005), Smirnova et al. (2010), Duyusen and Görkern (2011), McDonough (2014), He et al. (2015), Worsfold et al. (2019), Balaram et al. (2022) and Mnculwane (2022).

Attention is also drawn to the Geostandards Newsletter, which is currently called Geostandards and Geoanalytical Research and published by Wiley-Blackwell on behalf of the International Association of Geoanalysts.

## References

Note: All hyperlinks were checked on the 23rd of January 2022.

Agrawal, Y.C., McCave, I.N. & Riley, J.B., 2010. Laser diffraction size analysis. Chapter 9 In: J.P.M. Syvitski (Editor), Principles, Methods and Application of Particle Size Analysis. Cambridge University Press, 119−128; https://doi.org/10.1017/CBO9780511626142.012.

Alharbi, A. & El-Taher, A., 2016. Elemental analysis of basalt by instrumental neutron activation analysis and inductively coupled plasma mass spectrometer. Journal of Environment Science and Technology, 9(4), 335−339; https://doi.org/10.3923/jest.2016.335.339.

Andersson, M., Carlsson, M., Ladenberger, A., Morris, G., Sadeghi, M. & Uhlbäck, J., 2014.  Geokemisk atlas över Sverige - Geochemical Atlas of Sweden.  Geological Survey of Sweden, Uppsala, 208 pp.; http://resource.sgu.se/dokument/mineralnaring/Geokemisk%20atlas/text/geokemisk-atlas-over-sverige.pdf.

Asamoah, R.E., Amankwah, R.K. & Addai-Mensah, J. 2014. Cyanidation of Refractory Gold Ores: A Review. 3rd UMat Biennial International Mining and Mineral Conference, 204−212; https://doi.org/10.13140/2.1.4772.6407.

ASTM, 2017. ASTM D7928: Standard Test Method for Particle-Size Distribution (Gradation) of Fine-Grained Soils Using the Sedimentation (Hydrometer) Analysis. American Society for Testing Materials, 24 pp.; https:/doi.org/10.1520/D7928-16.

Baker, J., Waight, T. & Ulfbeck, D., 2002. Rapid and highly reproducible analysis of rare earth elements by multiple collector inductively coupled plasma mass spectrometry. Geochimica et Cosmochimica Acta, 66(20), 3635−3646; https://doi.org/10.1016/S0016-7037(02)00921-3.

Balaram, V., 1995. Developments and trends in inductively coupled plasma mass spectrometry and its influence on the recent advances in trace element analysis. Current Science, 69(8), 640−649; https://www.jstor.org/stable/24097257.

Balaram, V., 2018. Recent advances and trends in Inductively Coupled Plasma-mass spectrometry and applications. Current Trends in Mass Spectrometry, LCGC Supplements, Special Issues-05-01-2018, 16(2), 8−13, 38; https://www.chromatographyonline.com/view/recent-advances-and-trends-inductively-coupled-plasma-mass-spectrometry-and-applications.

Balaram, V., 2020. Current and emerging analytical techniques for geochemical and geochronological studies. Geological Journal, Special Issue Article, 1−60; https://doi.org/10.1002/gj.4005.

Balaram, V., Mathur, R., Banakar, V.K., Hein, J.R., Rao, C.R.M., Rao, T.G. & Dasaram, B., 2006. Determination of the platinum-group elements (PGE) and gold (Au) in manganese nodule reference samples by nickel sulfide fire-assay and Te coprecipitation with ICP-MS. Indian Journal of Marine Sciences, 35(1), 7−16.

Balaram, V., Mathur, R., Satyanarayanan, M., Sawant, S.S., Roy, P., Subramanyam, K.S.V. & Dasaram, B, 2012. A rapid method for the determination of gold in rocks, ores and other geological materials by F-AAS and GF-AAS after separation and preconcentration by DIBK extraction for prospecting studies. MAPAN – Journal of Metrology Society of India, 27(2), 87−95; https://doi.org/10.1007/s12647-012-0012-2.

Balaram, V., Rahaman, W. & Roy, P., 2022. Recent advances in MC-ICP-MS applications in Earth and environmental sciences: Challenges and solutions. Geosystems and Geoenvironment, 1(2), 100019; https://doi.org/10.1016/j.geogeo.2021.100019.

Bank, T., Roth, E., Tinker, P. & Granite, E., 2016. Analysis of rare earth elements in geologic samples using Inductively Coupled Plasma Mass Spectroscopy. US DOE Topical Report, DOE/NETL-2016/1794, 10 pp.; https://netl.doe.gov/sites/default/files/netl-file/Rare-Earth-Trace-Bulk-Elemental-Analysis-ICP-MS-Topical-Report-4-14-2016.pdf.

Barefoot, R.R., 1998. Determination of precious metals in geological materials by inductively coupled plasma mass spectrometry. Journal Analytical Atomic Spectrometry, 13, 1077−1084; https://doi.org/10.1039/A803216E.

Benedetti, M.F., De Kersabiec, A.M. & Bouleque, J., 1987. Determination of gold in twenty geochemical reference samples by flameless atomic absorption spectrometry. Geostandards Newsletter, 11(1), 127−129; https://doi.org/10.1111/j.1751-908X.1987.tb00016.x.

BGS, 1978. Shetland: Regional geochemical atlas. British Geological Survey, Keyworth, Nottingham, 37 pp.

BGS, 1991. Regional geochemistry of the East Grampians area. British Geological Survey, Keyworth, Nottingham, 94 pp.

BGS, 1992. Regional geochemistry of the Lake district and adjacent areas. British Geological Survey, Keyworth, Nottingham, 98 pp.

BGS, 1993. Regional geochemistry of southern Scotland and part of northern England. British Geological Survey, Keyworth, Nottingham, 96 pp.

BGS, 1999. Regional geochemistry of Wales and part of west-central England: Stream water. British Geological Survey, Keyworth, Nottingham, 110 pp.

BGS, 2000. Regional geochemistry of Wales and west-central England: Stream sediment and soil. British Geological Survey, Keyworth, Nottingham, 156 pp.

Birke, M., Rauch, U., Raschka, H., Wehner, H., Kringel, R., Gäbler, H.-E., Kriete, C., Siewers, U. & Kantor, W., 2006. Geochemischer Atlas Bundesrepublik Deutschland – Verteilung anorganischer und organischer Parameter in Oberflächenwässern und Bachsedimenten. Vorabexemplar, unpublished, 641 (in German).

Birke, M., Reimann, C., Demetriades, A., Rauch, U., Lorenz, H., Harazim, B. & Glatte, W., 2010. Determination of major and trace elements in European bottled mineral water — Analytical methods. In: M. Birke, A. Demetriades, B. De Vivo (Guest Editors), Mineral Waters of Europe. Special Issue, Journal of Geochemical Exploration, 107(3), 217−226; https://doi.org/10.1016/j.gexplo.2010.05.005.

Birke, M., Reimann, C. & Karl Fabian, K., 2014. Analytical Methods Used in the GEMAS Project. Chapter 5 In: C. Reimann, M. Birke, A. Demetriades, P. Filzmoser & P. O’Connor (Editors), Chemistry of Europe's agricultural soils – Part A: Methodology and interpretation of the GEMAS data set.  Geologisches Jahrbuch (Reihe B102), Schweizerbarth, Hannover, 41−46; https://www.schweizerbart.de/publications/detail/isbn/9783510968466.

Bounouira, H., Choukri, A., Elmoursli, R.C., Hakam, O.K. & Chakiri, S., 2007. Distribution of the rare earth elements in the sediments of the Bouregreg river (Morocco) using the instrumental neutron activation analysis (INAA). Journal of Applied Sciences and Environmental Management, 11(1), 57−60; https://doi.org/10.4314/jasem.v11i1.46834.

Bradshaw, N., Hall, E.F.H., & Sanderson, N.E., 1989. Communication. Inductively coupled plasma as an ion source for high resolution mass spectrometry. Journal of Analytical Atomic Spectrometry, 4, 801−803; https://doi.org/10.1039/JA9890400801.

Caritat, P. de & Cooper, M., 2011. National Geochemical Survey of Australia: the Geochemical Atlas of Australia. Geoscience Australia Record 2011/20, 557 pp. (2 Volumes);  http://www.ga.gov.au/metadata-gateway/metadata/record/gcat_71973.

Caritat, P. de, Reimann, C., Smith, D.B. & Wang, X., 2018. Chemical elements in the environment: Multi-element geochemical datasets from continental- to national-scale surveys on four continents. Applied Geochemistry, 89, 150−159; https://doi.org/10.1016/j.apgeochem.2017.11.010.

CEN, 2002. CEN Method 3052 – Microwave assisted acid digestion of siliceous and organically based matrices. CEN 13656, European Committee for Standardisation.

Chow, A. & Beamish, F.E., 1967. An experimental evaluation of neutron activation, wet assay and fire assay methods of determining gold in ores. Talanta, 14(2), 219−231; https://doi.org/10.1016/0039-9140(67)80181-4.

Coakley, J.P. & Syvitski, P.M., 2010. SediGraph technique. Chapter 10 In: J.P.M. Syvitski (Editor), Principles, Methods and Application of Particle Size Analysis. Cambridge University Press, 129−142; https://doi.org/10.1017/CBO9780511626142.013.

Coates, G.F. & Hulse, C.A., 1985. A comparison of four methods of size analysis of fine-grained sediments. New Zealand Journal of Geology and Geophysics, 28(2), 369−380; https://doi.org/10.1080/00288306.1985.10422234.

Darnley, A.G., Björklund, A., Bølviken, B., Gustavsson, N., Koval, P.V., Plant, J.A., Steenfelt, A., Tauchid, M., Xuejing, Xie., Garrett, R.G. & Hall, G.E.M., 1995. A Global Geochemical Database for Environmental and Resource Management. Recommendations for International Geochemical Mapping – Final Report of IGCP Project 259. Earth Science Report 19. UNESCO Publishing, Paris, 122 pp.; http://globalgeochemicalbaselines.eu.176-31-41-129.hs-servers.gr/datafiles/file/Blue_Book_GGD_IGCP259.pdf.

Demetriades, A., 2011. Understanding the quality of chemical data from the urban environment –  Part 2: Measurement uncertainty in the decision-making process. Chapter 6 In: C.C. Johnson, A. Demetriades, J. Locutura & R.T. Ottesen (Editors), Mapping the chemical environment of urban areas.  Wiley-Blackwell, John Wiley & Sons Ltd., Chichester, West Sussex, U.K., 77−98; https://doi.org/10.1002/9780470670071.ch6.

Demetriades, A., 2014. Basic considerations: Sampling, the key for a successful applied geochemical survey for mineral exploration and environmental purposes. Chapter 15.1 In: W. F. McDonough (volume Editor), Analytical geochemistry/Inorganic instrument analysis. In: H.D. Holland & K.K. Turekian (Executive Editors), Treatise on Geochemistry. Elsevier, Oxford, vol. 15, 1−31; https://doi.org/10.1016/B978-0-08-095975-7.01401-7.

Demetriades, A., Reimann, C. & Filzmoser, P., 2014. Evaluation of GEMAS project quality control results. Chapter 6 In: C. Reimann, M. Birke, A. Demetriades, P. Filzmoser & P. O’Connor (Editors), Chemistry of Europe's agricultural soils – Part A:  Methodology and interpretation of the GEMAS data set  Geologisches Jahrbuch (Reihe B102), Schweizerbarth, Hannover, 47−60; https://www.schweizerbart.de/publications/detail/isbn/9783510968466.

De Vos, W., Tarvainen, T., Salminen, R., Reeder, S., De Vivo, B., Demetriades, A., Pirc, S., Batista, M.J., Marsina, K., Ottesen, R.T., O’Connor, P.J., Bidovec, M., Lima, A., Siewers, U., Smith, B., Taylor, H., Shaw, R., Salpeteur, I., Gregorauskiene, V., Halamic, J., Slaninka, I., Lax, K., Gravesen, P., Birke, M., Breward, N., Ander, E.L., Jordan, G., Duris, M., Klein, P., Locutura, J., Bel-lan, A., Pasieczna, A., Lis, J., Mazreku, A., Gilucis, A., Heitzmann, P., Klaver, G. & Petersell, V., 2006. Geochemical Atlas of Europe. Part 2 - Interpretation of Geochemical Maps, Additional Tables, Figures, Maps, and Related Publications.  Geological Survey of Finland, Espoo, 692 pp.; http://weppi.gtk.fi/publ/foregsatlas/.

DIN, 1983. DIN 38406-5, Deutsche Einheitsverfahren zur Wasser-, Abwasser- und Schlammuntersuchung; Kationen (Gruppe E); Bestimmung des Ammonium-Stickstoffs (E 5). (German standard methods for the examination of water, waste water and sludge; cations (group E); determination of ammonia-nitrogen. Deutsches Institut für Normung e.V., Berlin.

DIN, 1988. DIN 38406-22, Edition, March 1988 - German standard methods for the examination of water, waste water and sludge Cations (Group E) Determination of Ag,Al,As,B,Ba,Be, Ca,Cd,Co,Cr, Cu,Fe,K,Li,Mg,Mn,Mo,Na,Ni,P,Pb,S,Sb,Se,Si,Sn,Ti,V,W,Zn and Zr by inductively coupled plasma atomic emission spectrometry (E 22), Deutsches Institut für Normung e.V., Berlin.

DIN, 1994. DIN 32645, Chemische Analytik - Nachweis-, Erfassungs- und Bestimmungsgrenze unter Wiederholbedingungen - Begriffe, Verfahren, Auswertung. (Chemical analysis - Decision limit, detection limit and determination limit under repeatability conditions - Terms, methods, evaluation). Deutsches Institut für Normung e.V., Berlin.

DIN, 1995. Wasserbeschaffenheit - Bestimmung der gelösten Anionen Fluorid, Chlorid, Nitrit, Orthophosphat, Bromid, Nitrat und Sulfat mittels Ionenchromatographie - Teil 1: Verfahren für gering belastete Wässer (ISO 10304-1:1992); Deutsche Fassung EN ISO 10304-1:1995 (Water quality - Determination of dissolved anions by liquid chromatography of ions - Part 1: Determination of bromide, chloride, fluoride, nitrate, nitrite, phosphate and sulfate (ISO 10304-1:1992; German version DIN EN ISO 10304-1:1995). Deutsches Institut für Normung e.V., Berlin.

DIN, 1996. DIN-Norm 38406, Teil 29. Analyse von 61 Elementen in Wasserproben mit ICP-MS.   Berlin (german industrial norm 38406, part 29, analyses of 61 elements in water samples by ICP-MS). Deutsches Institut für Normung e.V., Berlin.

DIN, 1998. DIN EN ISO 11885, Wasserbeschaffenheit - Bestimmung von ausgewählten Elementen durch induktiv gekoppelte Plasma-Atom-Emissionsspektrometrie (ICP-OES).  Deutsche Version EN ISO 11885:1997. (Water quality - Determination of 33 elements by inductively coupled plasma atomic emission spectroscopy (ISO 11885:1996; German version EN ISO 11885:1997). Deutsches Institut für Normung e.V., Berlin.

DIN, 2009. DIN EN ISO 11885, Edition, September 2009. Water quality - Determination of selected elements by inductively coupled plasma optical emission spectrometry (ICP-OES) (ISO 11885:2007). Deutsches Institut für Normung e.V., Berlin.

Dinnebier, R.E. & Billinge, S.J.L., 2008. Powder Diffraction: Theory and Practice. Royal Society of Chemistry, Cambridge, 582 pp.

Duyusen, E.G. & Görkern, A., 2011. Comparison of acid digestion techniques to determine heavy metals in sediment and soil samples. Gazi University Journal of Science, 24(1), 29−34; https://dergipark.org.tr/tr/download/article-file/83026.

EA, 1999. Expression of the uncertainty of measurement in calibration. European Accreditation, EA-4/02, rev.02, November 2021, 78 pp.; https://www.accredia.it/en/documento/ea-4-02-rev-02-expression-of-the-uncertainty-of-measurement-in-calibration/.

EA, 2003a. The selection and use of reference materials. European Accreditation, EA-4/14 INF, February 2003, 14 pp.; https://www.accredia.it/en/documento/ea-4-14-rev-00-the-selection-and-use-of-reference-materials/.

EA, 2003b. EA Guidelines on the expression of uncertainty in quantitative testing. European Accreditation, EA-4/16 rev. 00, December 2003, 16 pp.; https://www.accredia.it/en/documento/ea-416-rev-00-ea-guidelines-on-the-expression-of-uncertainty-in-quantitative-testing/.

EA, 2013. Evaluation of the uncertainty of measurement in calibration. European Accreditation, EA-4/02 M:2013, Revision 01, 75 pp.; https://www.kral-usa.com/fileadmin/data/files/services/European_Accreditation_EA-4-02.pdf.

Eggins, S.M., 2003. Laser ablation ICP-MS analysis of geological materials prepared as lithium borate glasses. Geostandards and Geoanalytical Research, 27(2), 147–162; https://doi.org/10.1111/j.1751-908X.2003.tb00642.x.

Ellison, S.L.R. & Williams, A., 2019. Eurachem/CITAC Guide: Metrological Traceability in Chemical Measurement – A guide to achieving comparable results in chemical measurement. Eurachem/CITAC, 2nd Edition, 38 pp.; https://www.eurachem.org/images/stories/Guides/pdf/ECTRC_2019_EN_P1.pdf.

Ellison, S.L.R., King, B., Rösslein, M., Salit, M. & Williams, A., 2003. EURACHEM/CITAC Guide: Traceability in Chemical Measurement – A guide to achieving comparable results in chemical measurement. Eurachem/CITAC, 37 pp.

Ely, J.C., Neal, C.R., O’Neil, J.A. Jr. & Jain, J.C., 1999. Quantifying the platinum group elements (PGEs) and gold in geological samples using cation exchange pretreatment and ultrasonic nebulization inductively coupled plasma mass spectrometry (USN-ICP-MS). Chemical geology, 157(3-4), 219−234; https://doi.org/10.1016/S0009-2541(98)00204-6.

Eurachem, 2002. The Selection and Use of Reference Materials. Eurachem, EEE/RM/062rev3, 10 pp.; https://www.eurachem.org/images/stories/Guides/pdf/EEE-RM-062rev3.pdf.

Fauth, H., Hindel, R., Siewers, U. & Zinner, J., 1985. Geochemischer Atlas der Bundesrepublik Deutschland. Schweizerbart'sche Verlagsbuchhandlung, Stuttgart, 79 pp.

Fedorowich, J.S., Richards, J.P., Jain, J.C., Kerrich, R. & Fan, J., 1993. A rapid method for REE and trace-element analysis using laser sampling ICP-MS on direct fusion whole-rock glasses. Chemical Geology, 106(3-4), 229−249; https://doi.org/10.1016/0009-2541(93)90029-I.

Figueiredo, A.M.G., Enzweiler, J., Sarkis, J.E.S., Jorge, A.P.S. & Shibuya, E.K., 2000. NAA and UV laser ablation ICP-MS for platinum-group elements and gold determination in NiS fire assay buttons: A comparison between two methods. Journal of Radioanalytical and Nuclear Chemistry, 244, 623−625; https://doi.org/10.1023/A:1006725618998.

Finch, C., Hall, G.E.M. & McConnell, J.W., 1992. The development and application of geochemical analyses. Newfoundland Dept. of Mines and Energy, Current Research 1992, Geological Survey Branch Report 92-1, 297−307.

Fletcher, W.K., 1981. Analytical methods in geochemical prospecting. Volume 1 In: G.J.S. Govett (Editor), Handbook of Exploration Geochemistry. Elsevier Scientific Publishing Co., Amsterdam, 255 pp.

Flight, D.M.A., Hall, G.E.M. & Simpson, P.R., 1994. Regional geochemical mapping of Pt, Pd and Au over an obducted ophiolite complex, Shetland Islands, northern Scotland. In: D.L. Buchanan (Editor), IAGOD international symposium on mineralization related to mafic and ultramafic rocks, Orleans, France, Sept. 1-3 1993. Transactions Institution of Mining and Metallurgy, Section B, Applied Earth Science, 103, 68−78.

Forbes, A.B., 2012. Approaches to evaluating measurement uncertainty. International Journal of Metrology and Quality Engineering, 3(2), 71−77; https://doi.org/10.1051/ijmqe/2012017.

Fuge, R. & Andrews, M.J., 1985. The automated photometric determination of total fluorine in mineral exploration. Journal of Geochemical Exploration, 23(3), 293−297; https://doi.org/10.1016/0375-6742(85)90032-9.

Fuge, R., Johnson, C.C. & Phillips, W.I., 1978. An automated method for the determination of iodine in geochemical samples. Chemical Geology, 23(1-4), 255-265; https://doi.org/10.1016/0009-2541(78)90081-5.

Gäbler, H.-E., 2002. Application of magnetic sector ICP-MS in geochemistry. Journal of Geochemical Exploration, 75(1-3), 1−15; https://doi.org/10.1016/S0375-6742(01)00197-2.

Garuti, G., Meloni, S. & Oddone, M., 2000. NAA of platinum group elements and gold in reference materials: A comparison of two methods. Journal of Radioanalytical and Nuclear Chemistry, 245, 17−23; https://doi.org/10.1023/A:1006709904766.

Gent, A.C. & Wilson, A.S., 1985. The determination of sulfur and chlorine in coals and oil shales using ion chromatography. Analytical Letters, 18(A6), 729−740; https://doi.org/10.1080/00032718508066171.

Giessmann, U. & Grebb, U., 1994. High resolution ICP-MS – a new concept for elemental mass spectrometry. Fresenius’ Journal of Analytical Chemistry, 350, 186−193; https://doi.org/10.1007/BF00322469.

Halamić, J. & Miko, S. (Editors), 2009. Geokemijski Atlas Republike Hrvatske – Geochemical Atlas of the Republic of Croatia. Croatian Geological Survey, Zagreb, 87 pp.

Hall, G.E.M., 1991. Analytical methods used in exploration geochemistry. In: J.M. Franklin, J.M. Duke, W.W. Shilts, W.B. Coker, P.W.B. Friske, Y.T. Marice, S.B. Ballantyne, C.E. Dunn, G.E.M. Hall & R.G. Garrett (Editors), Exploration Geochemistry Workshop.  Geological Survey of Canada Open File 2390; https://doi.org/10.4095/132388; http://ftp.maps.canada.ca/pub/nrcan_rncan/publications/ess_sst/132/132388/of_2390.pdf.

Hall, G.E.M. (Editor), 1992. Geoanalysis. Special Issue, Journal of Geochemical Exploration, 44, 349 pp.; https://www.sciencedirect.com/journal/journal-of-geochemical-exploration/vol/44/issue/1.

Hall, G.E.M. & Bonham-Carter, G.F., 1988. Review of methods to determine gold, platinum and palladium in production-oriented geochemical laboratories, with application of a statistical procedure to test for bias. Journal of Geochemical Exploration, 30(1-3), 255−286; https://doi.org/10.1016/0375-6742(88)90064-7.

Hall, G.E.M. & Pelchat, J.C., 1994. Analysis of geological materials for gold, platinum and palladium at low ppb levels by fire assay-ICP mass spectrometry. Chemical Geology, 115(1-2), 61−72; https://doi.org/10.1016/0009-2541(94)90145-7.

Hall, G.E.M. & Pelchat, J.C., 1997. Analysis of geological materials for bismuth, antimony, selenium and tellurium by continuous flow hydride generation inductively coupled plasma mass spectrometry, Part 1. Mutual hydride interferences. Journal of Analytical Atomic Spectrometry, 12(1), 97−102; https://doi.org/10.1039/A605398J.

Hall, G.E.M., MacLaurin, A.I. & Vaive, J., 1986. The analysis of geological materials for fluorine, chlorine and sulphur using pyrohydrolysis and ion chromatography. Journal of Geochemical Exploration, 26(2), 177−186; https://doi.org/10.1016/0375-6742(86)90066-X.

Hall, G.E.M., Vaive, J.E. & McConnell, J.W., 1995. Development and application of a sensitive and rapid analytical method to determine rare-earth elements in surface waters. Chemical Geology, 120, 91−109; https://doi.org/10.1016/0009-2541(94)00110-T.

Hall, G.E.M., Vaive, J.E. & Pelchat, J.C., 1996. Performance of ICP mass spectrometric methods used in the determination of trace elements in surface waters in hydrogeochemical surveys. Journal of Analytical Atomic Spectrometry, 11, 779−786;  https://doi.org/10.1039/ja9961100779.

Hann, S., Koellensperger, G., Kanitsar, K. & Stingeder, G., 2001. ICP-SFMS determination of palladium using IDMS in combination with on-line and off-line matrix separation. Journal of Analytical Atomic Spectrometry, 16(9), 1057−1063; https://doi.org/10.1039/B102574K.

He, Z., Huang, F., Yu, H., Xiao, Y., Wang, F., Li, Q., Xia, Y., Zhang, X., 2015. A flux-free fusion technique for rapid determination of major and trace elements in silicate rocks by LA-ICP-MS. Geostandards and Geoanalytical Research, 40(1), 5−27; https://doi.org/10.1111/j.1751-908X.2015.00352.x.

Hoffstetter, A., Troll, G. & Matthies, D., 1991. Determination of trace amounts of fluorine, boron, and chlorine from a single sodium carbonate fusion of small geological sample masses. Analyst, 116(1), 65−67; https://doi.org/10.1039/AN9911600065.

Hooper, P.R., 1964. Rapid analysis of rocks by X-ray fluorescence. Analytical Chemistry, 36(7), 1271−1276; https://doi.org/10.1021/ac60213a026.

Hooper, P.R., 1966. A comparison of sample preparation methods in use for XRF analysis of rocks. In: J. Buwalda (Editor), Proceedings of the 5th Conference on X-ray Analytical Methods. University of Swansea (Wales, UK), Phillips Scientific and Analytical Equipment Department (Eindhoven, Netherlands), 76−87.

Hooper, P.R., Reidel, S.P., Brown, J.C., Bush, J.H., Holden, G.S., Kieck, W.D., Robinette, M., Sundstrom, C.E. & Taylor, T.L., 1976. Major element analyses of Columbia River basalt, part I; Basalt Research Group. Washington State University Open File report, 19 pp.

Hu, Z. & Gao, S., 2008. Upper crustal abundances of trace elements: A revision and update. Chemical Geology, 253(3-4), 205−221; https://doi.org/10.1016/j.chemgeo.2008.05.010.

ISO, 1995. ISO 10694: Soil quality - Determination of organic and total carbon after dry combustion (elementary analysis). International Standards Organization, Geneva, Switzerland, 5 pp.

ISO, 1997. ISO 11843-1:1997/COR 1:2003: Capability of Detection – Part 1: Terms and Definitions – Technical Corrigendum 1, 1997-07 and 2003-11. International Standards Organization, Geneva, Switzerland, 10 pp.

ISO, 2017. BS EN ISO 17025: General requirements for the competence of testing and calibration laboratories. International Standards Organization, Geneva, Switzerland, 38 pp.; https://www.iso.org/obp/ui/#iso:std:iso-iec:17025:ed-3:v1:en.

ISO, 2018. ISO 11843-7: Capability of Detection – Part 7: Methodology based on stochastic properties of instrumental noise, 2018-09, 18 pp.

ISO, 2020a. ISO 5725-4: Parts 1 to 6: Accuracy (trueness and precision) of measurement methods and results. International Standards Organization, Geneva, Switzerland, 26 pp. (ISO 5725:1994 is replaced by ISO 5725-4:2020).

ISO, 2020b. ISO 11277: Soil quality - Determination of particle size distribution in mineral soil material - Method by sieving and sedimentation. International Organisation for Standardization, Geneva, Switzerland, 34 pp. (ISO 11277:2009 is replaced by ISO 11277:2020).

ISO, 2020c. ISO 13320: Particle size analysis - Laser diffraction methods. International Organisation for Standardization, Geneva, Switzerland, 59 pp. (ISO 13320:2009 is replaced by ISO 13320:2020).

Jackson, S.E., Fryer, B.J., Gosse, W., Healey, D.C., Longerich, H.P. & Strong, D.F., 1990. Determination of the precious metals in geological materials by inductively coupled plasma mass spectrometry (ICP-MS) with nickel sulphide fire-assay collection and tellurium coprecipitation. Chemical Geology, 83(1-2), 119−132; https://doi.org/10.1016/0009-2541(90)90144-V.

Janik, L.J., Forrester, S.T. & Rawson, A., 2009. The prediction of soil chemical and physical properties from mid-infrared spectroscopy and combined partial least-squares regression and neural networks (PLS-NN) analysis. Chemometrics and Intelligent Laboratory Systems, 97(2), 179−188; https://doi.org/10.1016/j.chemolab.2009.04.005.

Janik L., McLaughlin, M.J. & Forrester, S., 2011. Predicting particle size distribution in soils of the GEMAS sampling program using infrared spectroscopy and partial least squares regression. CSIRO Commercial-in-confidence report, 22 pp. Appendix 3 In: C. Reimann, A. Demetriades, M. Birke, O.A. Eggen, P. Filzmoser, P., C. Kriete & EuroGeoSurveys Geochemistry Expert Group (2012): The EuroGeoSurveys Geochemical Mapping of Agricultural and grazing land Soils project (GEMAS) – Evaluation of quality control results of particle size estimation by MIR prediction, Pb-isotope and MMI® extraction analyses and results of the GEMAS ring test for the standards Ap and Gr. NGU Report 2012.051, 136 pp. http://www.ngu.no/upload/Publikasjoner/Rapporter/2012/2012_051.pdf.

Janik, L., Forrester, S., Kirby, J.K., Soriano-Disla, J.M., Kirby, J.K., McLaughlin, M.J., Reimann, C. & The GEMAS Project Team, 2015a. GEMAS: Prediction of solid-solution partitioning coefficients (Kd) for cationic metals in soils using mid-infrared diffuse reflectance spectroscopy. Environmental Toxicology and Chemistry, 34(2), 224−234; https://doi.org/10.1002/etc.2736.

Janik, L., Forrester, S., Soriano-Disla, J.M., Kirby, J.K., McLaughlin, M.J., Reimann, C. & The GEMAS Project Team, 2015b. GEMAS: Prediction of solid-solution phase partitioning coefficients (Kd) for oxoanions and boric acid in soils using mid-infrared diffuse reflectance spectroscopy.  Environmental Toxicology and Chemistry, 34(2), 235−246; https://doi.org/10.1002/etc.2821.

Jarvis, K.E., Gray, A.L. & Houk, R.S., 1992. Handbook of Inductively Coupled Plasma Mass Spectrometry. Blackie, Glasgow, 380 pp.

Jenkins, R. & Snyder, R.I., 1996. Introduction to X-ray Powder Diffractometry. Chemical Analysis, vol. 138. John Wiley & Sons, Inc., 403 pp.; https://doi.org/10.1002/9781118520994.

Johnson, D.M., Hooper, P.R. & Conrey, R.M., 1999. XRF analysis of rocks and minerals for major and trace elements on a single low dilution Li-tetraborate fused bead. Advances in X-ray Analysis, 41, 843−867.

Jordan, J.R., Fryer, G.E. & Hemmen, E.H., 1971. Size analysis of silt and clay by hydrophotometer. Journal of Sedimentary Research, 47(2), 931−932; https://doi.org/10.1306/74D722AC-2B21-11D7-8648000102C1865D.

Kadünas, V., Budavičius, R., Gregorauskiene, V., Katinas, V., Kliaugienė, E., Radzevičius, A. & Taraškevičius, R., 1999. Lietuvos Geocheminis Atlasas – Geochemical Atlas of Lithuania.  Geological Survey of Lithuania, Vilnius, 161 pp.

Keays, R.R. & Scott, R.B., 1976. Precious metals in ocean ridge basalts: implications for basalts as source rocks for gold mineralization. Economy Geology, 71(4), 705−720; https://doi.org/10.2113/gsecongeo.71.4.705.

Kennedy, W.T., Hubbard, W.B. & Tarter J.G., 1983. Rapid analysis of fluorine in geological samples with ion chromatographic detection. Analytical Letters, 16(15), 1133−1148; https://doi.org/10.1080/00032718308077151.

Kent, A.J.R., Jacobsen, B., Peate, D.W., Waight, T.E. & Baker, J.A., 2007. Isotope dilution MC-ICP-MS rare earth element analysis of geochemical reference materials NIST SRM 610, NIST SRM 612, NIST SRM 614, BHVO-2G, BHVO-2, BCR-2G, JB-2, WS-E, W-2, AGV-1 and AGV-2. Geostandards and Geoanalytical Research 28(3), 417−429; https://doi.org/10.1111/j.1751-908X.2004.tb00760.x.

Koljonen, T. (Editor), 1992. The Geochemical Atlas of Finland, Part 2: Till. Geological Survey of Finland, Espoo, 218 pp.

Kontas, E., Niskavaara, H. & Virtasalo, J., 1990. Gold, palladium and tellurium in South African, Chinese and Japanese geological reference samples. Geostandards Newsletter, 14(3), 477−478; https://doi.org/10.1111/j.1751-908X.1990.tb00083.x.

Krachler, M., Mohl, C., Emons, H. & Shotyk, W., 2002. Journal of Analytical Atomic Spectrometry, 17(8), 844−851; https://doi.org/10.1039/B200780K.

Kurosawa, M., Shima, K., Ishii, S. & Sasa, K., 2007. Trace element analysis of fused whole-rock glasses by laser ablation-ICP-MS and PIXE. Geostandards and Geoanalytical Research, 30(1), 17−30; https://doi.org/10.1111/j.1751-908X.2006.tb00908.x.

Lahermo, P., Ilmasti, M., Juntunen, R. & Taka, M., 1990.  Suomen Geokemian Atlas, osa 1 – Suomen pohjavesien hydrogeokemiallinen kartoitus − The Geochemical Atlas of Finland, Part 1: The hydrogeochemical mapping of Finnish groundwater. Geological Survey of Finland, Espoo, 66 pp.

Lakin, H.W., Ward, F.N. & Almond, H., 1952. Compilation of field methods used in geochemical prospecting by the U.S. Geological Survey. United States Department of the Interior, Washington, D.C., United States Geological Survey, Circular 161, 34 pp.; https://doi.org/10.3133/cir161.

Langenauer, M., Krähenbühl, U., Furrer, V. & Wyttenbach, A., 1992. Determination of fluorine, chlorine, bromine and iodine in seven geochemical reference samples. Geostandards Newsletter, 16(1), 41−45; https://doi.org/10.1111/j.1751-908X.1992.tb00485.x.

Lavina, B., Dera, P. & Downs, R.T., 2014. Modern X-ray Diffraction Methods in Mineralogy and Geosciences. Review in Mineralogy & Geochemistry, 78(1), 1−31; https://doi.org/10.2138/rmg.2014.78.1.

Levinson, A.A., 1974. Introduction to Exploration Geochemistry. Applied Publishing Ltd., Wilmette, Illinois, USA, 614 pp.

Levinson, A.A., 1980. Introduction to exploration geochemistry. Applied Publishing Ltd., Wilmette, Illinois, USA, 924 pp.

Lis, J. & Pasieczna, A., 1995. Geochemical Atlas of Poland. Polish Geological Institute, Warsaw, 36 pp., 7 Tables, 76 maps.

Liu, Y., Hu, Z., Li, M. & Gao, S., 2013. Applications of LA-ICP-MS in the elemental analyses of geological samples. Chinese Science Bulletin, 58, 3863–3878; https://doi.org/10.1007/s11434-013-5901-4.

Magnusson, B. & Örnemark, U., 2014. Eurachem Guide: The Fitness for Purpose of Analytical Methods – A Laboratory Guide to Method Validation and Related Topics, Second edition, 70 pp.; https://www.eurachem.org/images/stories/Guides/pdf/MV_guide_2nd_ed_EN.pdf.

Marsden, J.O. & House, I.C., 2006. The Chemistry of Gold Extraction. Society for Mining, Metallurgy, and Exploration, Inc. (SME), Englewood, Colorado, 651 pp.

McDonough, W.F., 2014. Analytical Geochemistry/Inorganic Instrument Analysis. Volume 15 In: H.D. Holland & K.K. Turekian (Executive Editors), Treatise on Geochemistry. Elsevier, Second edition, 454 pp.

Mnculwane, H.T., 2022. Rare earth elements determination by inductively coupled plasma mass spectrometry after alkaline fusion preparation. Analytica, 3(1), 135–143; https://doi.org/10.3390/analytica3010010.

Montaser, A. (Editor), 1998. Inductively Coupled Plasma Mass Spectrometry. Wiley-VCH, New York 1008 pp.

Morcelli, C.P.R., Figueiredo, A.M.G., Enzweiler, J., Sarkis, J.E.S., Jorge, A.P.S. & Kakazu, M., 2004. Determination of platinum-group elements in geological reference materials by high resolution-ICP-MS after nickel sulfide fire-assay collection and Te co-precipitation. Geostandards and Geoanalytical Research, 28(2), 305−310; https://doi.org/10.1111/j.1751-908X.2004.tb00745.x.

Müller, H.-W., Dohrmann, R., Klosa, D., Rehder, S. & Eckelmann, W., 2009. Comparison of two procedures for particle-size analysis: Köhn pipette and X-ray granulometry. Journal of Plant Nutrition and Soil Science, 172(2), 172−179; https://doi.org/10.1002/jpln.200800065.

Nesbitt, B.E., St Louis, R.M. & Muehlenbachs K., 1987. Distribution of gold in altered basalts of DSDP hole 504B. Canadian Journal of Earth Sciences, 24(2), 201−209; https://doi.org/10.1139/e87-022.

Odegard, M., Dundas, S.H., Flem, B. & Grimstvedt, V., 1998. Application of a douple-focusing magnetic sector inductively coupled plasma mass spectrometer with laser ablation fort the bulk analysis of rare earth elements in rocks fused with Li2B4O7. Fresenius Journal of Analytical Chemistry, 362, 477−482.

Olesik, J.W., 2000. Inductively Coupled Plasma Mass Spectrometry. In: C.M. Barshick, D.C. Duckworth & D.H. Smith (Editors), Inorganic Mass Spectrometry, Fundamentals and Applications. Marcel Dekker, New York, 67-158.

Ottesen, R.T., Bogen, J., Bølviken, B., Volden, T. & Haugland, T., 2000. Kjemisk sammensetning av flomsedimenter - Geochemical Atlas of Norway, Part 1: Chemical Composition of Overbank Sediments. Norges geologiske undersøkelse, NGU, Geological Survey of Norway, Trondheim, 140 pp.

Ottesen, R.T., Bogen, J., Finne, T.E., Andersson, M., Dallmann, W.K., Eggen, O.A., Jartun, M., Lundkvist, Q., Pedersen, H.R. & Volden, T., 2010. Geochemical Atlas of Norway, Part 2:  Geochemical Atlas of Spittsbergen – Chemical Composition of Overbank Sediments. Geological Survey of Norway, Trondheim, 160 pp.

Park, J.-W., Hu, Z., Gao, S., Campbell, I.H. & Gong, H., 2012. Platinum group element abundances in the upper continental crust revisited – New constraints from analyses of Chinese loess. Geochimica et Cosmochimica Acta, 93, 63−76; https://doi.org/10.1016/j.gca.2012.06.026.

Pecharsky, V.K. & Zavalij, P.Y., 2009. Fundamentals of Powder Diffraction and Structural Characterisation of Materials, Springer, New York, Second edition, 703 pp.

Pitcairn, I.K., Warwick, P.E., Milton, J.A., & Teagle, D.A.H., 2006. Method for ultra-low level analysis of gold in rocks. Analytical Chemistry, 78(4), 1280−1285; https://doi.org/10.1021/ac051861z.

Polakowski, C., Ryżak, M., Sochan, A., Beczek, M., Mazur, R. & Bieganowski, A., 2021. Particle size distribution of various soil materials measured by laser diffraction – The problem of reproducibility. Minerals, 11, 465, 10 pp.; https://doi.org/10.3390/min11050465.

Ramesh, S.L., Raju, P.V.S., Anjaiah, K.V., Mathur, R., Rao, T.G., Dasaram, B., Charan, N.S., Rao, D.V.S., Sarma, D.S., Mekala, R.M. & Balaram, V., 2001. Determination of gold in rocks, ores and other geological materials by atomic absorption techniques. Atomic Spectroscopy, 22(1), 263−269.

Ravisankar, R., Manikandan, E., Dheenathayalu, M., Rao, B., Seshadreesan, N.P. & Nair, K.G.M., 2006. Determination and distribution of rare earth elements in beach rock samples using instrumental neutron activation analysis (INAA). Nuclear Instruments and Methods in Physics Research, Section B: Beam Interactions with Materials and Atoms, 251(2), 496−500; https://doi.org/10.1016/j.nimb.2006.07.021.

Reddi, G.S. & Rao, C.R.M., 1999. Analytical techniques for the determination of precious metals in geological and related materials. Analyst, 124, 1531−1541; https://doi.org/10.1039/A904839A.

Reimann, C. & Birke, M. (Editors), 2010. Geochemistry of European Bottled Water. Borntraeger Science Publishers, Stuttgart, 268 pp.; http://www.schweizerbart.de/publications/detail/artno/001201002#.

Reimann, C., Demetriades, A., Eggen, O.A., Filzmoser, P. & The EuroGeoSurveys Geochemistry Expert Group, 2009. The EuroGeoSurveys geochemical mapping of agricultural and grazing land soils project (GEMAS) – Evaluation of quality control results of aqua regia extraction analysis.  Geological Survey of Norway, Trondheim, NGU report 2009.049, 94 pp.; http://www.ngu.no/upload/Publikasjoner/Rapporter/2009/2009_049.pdf.

Reimann, C., Demetriades, A., Eggen, O.A., Peter Filzmoser, P. & the EuroGeoSurveys Geochemistry Expert Group, 2011. The EuroGeoSurveys GEochemical Mapping of Agricultural and grazing land Soils project (GEMAS) – Evaluation of quality control results of total C and S, total organic carbon (TOC), cation exchange capacity (CEC), XRF, pH, and particle size distribution (PSD) analysis. Geological Survey of Norway, Trondheim, NGU Report 2011.043, 90 pp.; http://www.ngu.no/upload/Publikasjoner/Rapporter/2011/2011_043.pdf.

Reimann, C., Demetriades, A., Birke, M., Eggen, O. A., Filzmoser, P., Kriete, C. & EuroGeoSurveys Geochemistry Expert Group, 2012. The EuroGeoSurveys Geochemical Mapping of Agricultural and grazing land Soils project (GEMAS) – Evaluation of quality control results of particle size estimation by MIR prediction, Pb-isotope and MMI® extraction analyses and results of the GEMAS ring test for the standards Ap and Gr. Geological Survey of Norway, Trondheim, NGU report 2012.051, 136 pp.; http://www.ngu.no/upload/Publikasjoner/Rapporter/2012/2012_051.pdf.

Rice, T.D., 1988. Determination of fluorine and chlorine in geological materials by induction furnace pyrohydrolysis and standard-addition ion-selective electrode measurement. Talanta, 35(3), 173−178: https://doi.org/10.1016/0039-9140(88)80060-2.

Ripley, E.M., Li, C., Moore, C.H., Elswick, E.R., Maynard, J.B., Paul, R.L., Sylvester, P., Seo, J.H., & Shimizu, N., 2011. Analytical methods for sulfur determination in glasses, rocks, minerals and fluid inclusions. Reviews in Mineralogy and Geochemistry, 73(1), 9−39; https://doi.org/10.2138/rmg.2011.73.2.

Rudnick, R.L. & Gao, S., 2003. Composition of the continental crust. In: H.D. Holland & K.K. Turekian (Editors-in-chief), Treatise on Geochemistry; Volume 3, R.L. Rudnick (Editor), The Crust.  Elsevier-Pergamon, Oxford, 1−64; https://doi.org/10.1016/B0-08-043751-6/03016-4.

Salminen, R., Batista, M.J., Bidovec, M., Demetriades, A., De Vivo, B., De Vos, W., Duris, M., Gilucis, A., Gregorauskiene, V., Halamic, J., Heitzmann, P., Lima, A., Jordan, G., Klaver, G., Klein, P., Lis, J., Locutura, J., Marsina, K., Mazreku, A., O’Connor, P.J., Olsson, S.Å., Ottesen, R.T., Petersell, V., Plant, J.A., Reeder, S., Salpeteur, I., Sandström, H., Siewers, U., Steenfelt, A. & Tarvainen, T., 2005.  FOREGS Geochemical Atlas of Europe, Part 1: Background Information, Methodology and Maps.  Geological Survey of Finland, Espoo, 526 pp.; http://weppi.gtk.fi/publ/foregsatlas/.

Sandström, H., Reeder, S., Bartha, A., Birke, M., Berge, F., Davidsen, B., Grimstvedt, A., Hagel-Brunnström, M-L., Kantor, W., Kallio, E., Klaver, G., Lucivjansky, P., Mackovych, D., Mjartanova, H., van Os, B., Paslawski, P., Popiolek, E., Siewers, U., Varga-Barna, Zs., van Vilsteren, E. & Ødegård, M., 2005. Sample preparation and analysis. In: R. Salminen (Chief-editor), M.J. Batista, M. Bidovec, A. Demetriades, B. De Vivo, W. De Vos, M. Duris, A. Gilucis, V. Gregorauskiene, J. Halamic, P. Heitzmann, A. Lima, G. Jordan, G. Klaver, P. Klein, J. Lis, J. Locutura, K. Marsina, A. Mazreku, P.J. O’Connor, S.Å. Olsson, R.T. Ottesen, V. Petersell, J.A. Plant, S. Reeder, I. Salpeteur, H. Sandström, U. Siewers, A. Steenfelt, & T. Tarvainen, FOREGS Geochemical Atlas of Europe, Part 1: Background Information, Methodology and Maps. Geological Survey of Finland, Espoo, 81−94; http://weppi.gtk.fi/publ/foregsatlas/articles/Analysis.pdf.

Satyanarayanan, M., Balaram, V., Sawant, S.S., Subramanyam, K.S.V., Krishna, G.V., Dasaram, B. & Manikyamba, C., 2018. Rapid determination of REEs, PGEs and other trace elements in geological and environmental materials by high resolution inductively coupled plasma mass spectrometry.. Atomic Spectroscopy, 39(1), 1−15; https://doi.org/10.46770/AS.2018.01.001.

Shazali, I., Van’t Dack, L. & Gijbels, R., 1987. Determination of precious metals in ores and rocks by thermal neutron activation/gamma-spectrometry after preconcentraion by nickel sulfide fire assay and coprecipitation with tellurium. Analytica Chimica Acta, 196, 49−58; https://doi.org/10.1016/S0003-2670(00)83069-X.

Singh, Th.D., Manikyamba, C., Subramanyam, K.S.V., Ganguly, S., Khelen, A.C. & Reddy, N.R., 2018. Mantle heterogeneity, plume-lithosphere interaction at rift-controlled ocean-continent transition zone: Evidence from trace-PGE geochemistry of Vempalle flows, Cuddapah Basin, India. Geoscience Frontiers, 9(6), 1809−1827; https://doi.org/10.1016/j.gsf.2017.12.013.

Smirnova, E.V., Flem, B., Anchutina, E.A., Mysovskaya, I.N., Lozhkin, V.I. & Petrov, L.L., 2010. Determination of REE, Y, Nb, Zr, Hf, Ta, Th and U in Geological Reference Materials LSHC-1 and Amf-1 by Solution and Laser Ablation ICP-MS. Geostandards and Geoanalytical Research, 34, 49−65; https://doi.org/10.1111/j.1751-908X.2009.00009.x.

Smith, D.B., Cannon, W.F., Woodruff, L.G. & Ellefsen, K.J., 2014. Geochemical and Mineralogical Maps for Soils of the Conterminous United States. United States Geological Survey Open-File Report 2014−1082, 386 pp.; https://pubs.usgs.gov/of/2014/1082/.

Soriano-Disla, J.M., Janik, L.J., Viscarra Rossell, R.A., Macdonald, L.M. & McLaughlin, M.J., 2014. The performance of visible, near-, and mid-infrared reflectance spectroscopy for prediction of soil physical, chemical, and biological properties. Applied Spectroscopy Reviews, 49(2), 139–186; https://doi.org/10.1080/05704928.2013.811081.

Stanton, R.E., 1966. Rapid methods of trace analysis for geochemical application. Edward Arnold (Publishers), Ltd., London, 96 pp.

Strnad, L., Mihaljevic, M. & Sebek, O., 2007. Laser ablation and solution ICP-MS determination of rare earth elements in USGS BIR-1G, BHVO-2G and BCR-2G glass reference materials. Geostandards and Geoanalytical Research, 29(3), 303−314; https://doi.org/10.1111/j.1751-908X.2005.tb00902.x.

Takeda, A., Yamasaki, S., Tsukada, H., Takaku, Y., Hisamatsu, S. & Tsuchiya, N., 2011. Determination of total contents of bromine, iodine and several trace elements in soil by polarizing energy-dispersive X-ray fluorescence spectrometry. Soil Science and Plant Nutrition, 57(1), 19-28; https://doi.org/10.1080/00380768.2010.548313.

Tanaka, K., Takahashi, Y. & Shimizu, H., 2007. Determination of rare earth element in carbonate using laser-ablation inductively-coupled plasma mass spectrometry: An examination of the influence of the matrix on laser-ablation inductively-coupled plasma mass spectrometry analysis. Analytica Chimica Acta, 583(2), 303–309; https://doi.org/10.1016/j.aca.2006.10.023.

Terashima, S., 1988. Determination of total carbon and sulfur in fifty-two geochemical reference samples by combustion and infrared absorption spectrometry. Geostandards Newsletter, 12(1), 249−252; https://doi.org/10.1111/j.1751-908X.1988.tb00052.x.

Tewari, R.K., Tarsekar, V.K. & Lokhande, M.B., 1990. Determination of gold in geological materials by flame atomic absorption spectrophotometry at ppm and sub-ppm levels by coprecipitation with mercury. Atomic Spectroscopy, 11(4), 1125-1127.

Thalmann, F., 1984. Geochemischer Atlas der Republik Österreich 1:1 000 000 (Böhmische Masse und Zentralzone der Ostalpen), Geologische Bundesanstalt, Wien.

Thomas, C.L., Hernandez-Allica, J., Dunham, S.J., McGrath, S.P. & Haefele, S.M., 2021. A comparison of soil texture measurements using mid-infrared spectroscopy (MIRS) and laser diffraction analysis (LDA) in diverse soils. Nature Research, Scientific Reports, 11:16, 12 pp.; https://doi.org/10.1038/s41598-020-79618-y.

Thompson, M. & Howarth, R.J., 1976. Duplicate analysis in geochemical practice. Part 1. Theoretical approach and estimation of analytical reproducibility. Analyst, 101, 690–698; http://dx.doi.org/10.1039/AN9760100690.

Thompson, M. & Howarth, R.J., 1978. A new approach to the estimation of analytical precision.  Journal of Geochemical Exploration, 9, 23–30; https://doi.org/10.1016/0375-6742(78)90035-3.

Townshend, A. & Poole, C.F. (Editors), 2005. Encyclopedia of Analytical Science. Second edition (10 Volumes set). Elsevier, 5000 pp.

Ufer, K., Stanjek, H., Roth, G., Dohrmann, R., Kleeberg, R. & Kaufhold, S., 2008a. Quantitative phase analysis of bentonites by the Rietveld method. Clays and Clay Minerals, 56(2), 272−282; https://doi.org/10.1346/CCMN.2008.0560210.

Ufer, K., Kleeberg, R., Bergmann, J., Curtius, H. & Dorhmann, R., 2008b. Refining real structure parameters of disordered layer structures within the Rietveld method. Zeitschrift für Kristallographie Suoolements, 27, 151−158; https://doi.org/10.1524/zksu.2008.0020.

Ufer, K., Kleeberg, R., Bergmann, J. & Dorhmann, R., 2012a. Rietveld refinement of disordered illite-smectite mixed layered structures by recursive algorithm. I: One-dimensional patterns. Clay and Clay Minerals, 60(5), 507−534; https://doi.org/10.1346/CCMN.2012.0600507.

Ufer, K., Kleeberg, R., Bergmann, J. & Dorhmann, R., 2012b. Rietveld refinement of disordered illite-smectite mixed layered minerals with a recursive algorithm. II: Powder pattern refinement and quantitative phase analysis. Clay and Clay Minerals, 60(5), 535−552; https://doi.org/10.1346/CCMN.2012.0600508.

Ullman, W.J. & Tisue, G.T., 1983. Determination of bromine and iodine in U.S.G.S. standard marine mud (MAG-1). Geostandards Newsletter, 7(2), 289−290; https://doi.org/10.1111/j.1751-908X.1983.tb00385.x.

USEPA, 1996. Method 3052 − Microwave assisted acid digestion of siliceous and organically based matrices. United States Environmental Protection Agency, 20 pp.; https://www.epa.gov/sites/default/files/2015-12/documents/3052.pdf.

USEPA, 1997. EPA-NERL: 415.1: Total Organic Carbon. Official Method Name: Organic Carbon, Total (Combustion or Oxidation). National Environmental Methods Index, United States Environmental Protection Agency, 3 pp.; https://www.nemi.gov/methods/method_summary/5403/.

USEPA, 2004. Method 9060A: Total Organic Carbon. United States Environmental Protection Agency, 5 pp.; https://www.epa.gov/sites/default/files/2015-12/documents/9060a.pdf.

USEPA, 2007. Method 3051A – Microwave assisted acid digestion of sediment, sludges, soils and oils. Revision 1, United States Environmental Protection Agency, 30 pp; https://www.epa.gov/sites/default/files/2015-12/documents/3051a.pdf.

Vanloon, J.C. & Barefoot, R.R., 1991. Determination of the Precious Metals – Selected Instrumental Methods. Wiley, Chichester, 276 pp.; https://doi.org/10.1002/xrs.1300210412.

Viscarra Rossell, R.A., Walvoort, D.J.J., McBratney, A.B., Janik, L.J. & Skjemstad, J.O., 2006. Visible, near infrared, mid-infrared or combined diffuse reflectance spectroscopy for simultaneous assessment of various soil properties. Geoderma, 131(1-2), 59–75; https://doi.org/10.1016/j.geoderma.2005.03.007.

Ward, F.N., Lakin, H.W., Canney, F.C. & others, 1963. Analytical methods used in geochemical exploration by the U.S. Geological Survey.  U.S. Geological Survey Bulletin 1152, United States Government Printing Office, Washington, 100 pp.; https://pubs.usgs.gov/bul/1152/report.pdf.

Waseda, Y., Matsubara, E. & Shinoda, K., 2011. X-ray Diffraction Crystallography. Springer-Verlag, Berlin-Heidelberg, 310 pp.; https://link.springer.com/book/10.1007/978-3-642-16635-8.

Webb, J.S., Thornton, I., Thompson, M., Howarth, R.J. & Lowenstein, P.L., 1978. The Wolfson geochemical atlas of England and Wales. Imperial College of Science and Technology. Applied Geochemistry Research Group. Clarendon Press, Oxford, 69 pp.

Wedepohl, K.H., 1969–1978. Handbook of Geochemistry, Vol. I–II/1–5. Springer Verlag, Berlin.

Wedepohl, K.H., 1995. The composition of the continental crust. Geochimica et Cosmochimica Acta, 59(7), 1217-1232; https://doi.org/10.1016/0016-7037(95)00038-2.

Welday, E.E., Baird, A.K., McIntyre, D.B. & Madlem, K.W., 1964. Silicate sample preparation for light element analyses by X-ray spectrography. American Mineralogist, 49(7-8), 889−903.

Wisconsin Department of Natural Resources, 1996. Analytical Detection Limit Guidance & Laboratory Guide for Determining Method Detection Limits. Laboratory Certification Program, 29, 33 pp.; http://www.iatl.com/content/file/LOD%20Guidance%20Document.pdf.

Worsfold, P., Townshend, A., Poole, C. & Miro, M. (Editors), 2019. Encyclopedia of Analytical Science. 3rd Edition (10 Volumes set), Elsevier.

Xie, X., Yan, M. & Xuejing, X., 1985. Provisional Geochemical Atlas of Northwestern Jiangchi, Geological Publishing House, Beijing, 46 pp.

Xie, X., Yan, M., Li, L. & Shen, H., 2007. Usable values for Chinese standard reference samples of stream sediments, soils, and Rocks: GSD 9-12, GSS1-8 AND GSR 1-6. Geostandards and Geoanalytical Research 9(2), 277−280; https://doi.org/10.1111/j.1751-908X.1985.tb00458.x.

Xie, X., Tianxiang, R. & Huanzhen, S., 2012. Geochemical Atlas of China. Zhongguo Diqiu Huaxue Tuji, 135 pp.

Yamasaki, T. & Yamashita, K., 2016. Whole rock multiple trace element analyses using fused glass bead by laser ablation-ICP-MS. Bulletin of the Geological Survey of Japan, 67(1), 27−40; https://www.gsj.jp/data/bulletin/67_01_03.pdf.

Yan, M., Wang, C., Cao, Q., Gu, T. & Chi, Q., 1995. Eleven gold geochemical reference samples (GAU 8-18). Geostandards Newsletter, 19(2), 125−133; https://doi.org/10.1111/j.1751-908X.1995.tb00157.x.

Yu, Z., Norman, M.D. & Robinson, P., 2003. Major and trace element analysis of silicate rocks by XRF and Laser Ablation ICP-MS using lithium borate fused glasses: Matrix effects, instrument response and results for international reference materials. Geostandards and Geoanalytical Research, 27(1), 67−89; https://doi.org/10.1111/j.1751-908X.2003.tb00713.x.

Yuchi, A., Yanai, N., Wada, H., & Nakagawa, G., 1988. Tiron as a making reagent for aluminium in the determination of fluoride with an ion-selective electrode: equilibrium studies and application. Analyst, 113, 1405−1408; https://doi.org/10.1039/AN9881301405.

Blank back page